Resolution in Gel Filtration:

Final resolution, the degree of separation between peaks, is influenced by many factors:

the ratio of sample volume to column volume, flow rate, column dimensions, particle size particle size distribution, packing density, porosity of the particle and viscosity of the mobile phase. The success of gel filtration depends primarily on choosing conditions that give sufficient selectivity and counteract peak broadening effects during the separation.

Resolution (Rs) is defined by the following expression:

Vr2 – Vr1

Rs =

(W1 + W2)

2

Vr1 and Vr2 are the elution volumes for two adjacent peaks measured at the center of the peak. W1 and W2 are the respective peak widths.

(Vr2 – Vr1) represents the distance between the peaks and 1/2 (W1 + W2) the mean peak width of the two peaks as shown in Figure.

Resolution is a function of the selectivity of the medium and the efficiency of that medium to produce narrow peaks (minimal peak broadening) as illustrated in down figure. The efficiency of a packed column defines its ability to produce narrow symmetrical peaks during elution. Efficiency is particularly important in gel filtration in which separation takes place as only a single column volume of buffer passes through the column.

Efficiency is defined in terms of theoretical plates per meter (N):

N = 5.54 (Ve/W1/2)2 × 1000/L

Where:

Ve = peak elution (retention) volume

W1/2 = peak width at half peak height

L = bed height (mm)

Ve and W1/2 are in same units

Satisfactory Separation

Well resolved, symmetrical peaks.

Poor Resolution

Review factors affecting resolution including media selection, particle size, sample volume: column volume and flow rate.

Leading peaks

Asymmetric peaks: sample elutes before void volume indicates channelling in column bed.Leading peaks can also be due to overpacking of column (packed at too high pressure or flow rate). Column may need to be repacked.

Tailing peaks

Asymmetric peaks: sample application uneven.Check top of column if possible. Ensure medium is evenly packed and that sample is applied without disturbing the packed bed. Tailing peaks can also be due to underpacking of column (packed at too low pressure or flow rate).

Late elution

Peaks seen after one column volume of buffer has passed through the column. Always include a wash step between runs to ensure removal of late eluting molecules.Molecules may be binding non-specifically to gel filtration medium. If the interaction is ionic in nature, increasing the concentration of sodium chloride (up to 150 mM) may help. If the interaction is hydrophobic in nature, reducing salt concentration, increasing pH or adding a detergent or organic solvent may help.

Molecular Weight Determination and Molecular Weight

Distribution Analysis:

1.If using a self-packed column, prepare a fresh, filtered solution of            Blue Dextran 2000 (1.0 mg/ml) in the running buffer. Apply Blue Dextran to the column, using a volume <2% of the total column volume (Vt) to determine the void volume (Vo).

2.Dissolve the selected calibration proteins in the running buffer (at concentrations recommended by the manufacturer). Allow a few minutes for dissolution, stirring gently. Do not heat or mix vigorously.If necessary, filter the calibration solution.

3.Apply the calibration solution to the column, in a volume <2% of the total column volume (Vt).

4.Determine the elution volumes (Ve) for the standards by measuring the volume of the eluent from the point of application to the centre of the elution peak.

5.Calculate the Kav values for the standards and prepare acalibration curve of Kav versus the logarithm of their molecular weights, as follows:

Ve – Vo

Kav =

Vt – Vo

where Ve = elution volume for the protein

Vo = column void volume = elution volume for Blue Dextran 2000

Vt = total bed volume

Plot the Kav value for each protein standard (on the linear scale) against the corresponding molecular weight (on the logarithmic scale). Draw the straight line which best fits the points on the graph. Alternatively, use a statistics package to calculate the regression line.

6.Apply the sample in a volume <2% of the total column volume (Vt) and determine the elution volume (Ve) of the molecule of interest.

7.Calculate the corresponding Kav for the component of interest and determine its molecular weight from the calibration curve.

Fig:Gel Filtration of Blue Dexran 2000 using Sephadex G-75

Fig.:Gel Filtration of Standard Protein was Carried out Using Sephadex G-75 ..

Table: Values Volume and Kav for Standard Proteins.

| Standard Protein | Conc. (mg/ml) | M.wt. ( KD ) | ( Ve ) ml | Kav |
| --- | --- | --- | --- | --- |
| Albumin | 2 | 67 | 48 | 0.2898 |
| Ova albumin | 2 | 43 | 65 | 0.4968 |
| Lactatedehydrogenase | 2 | 36 | 71.5 | 0.5796 |
| α-Chymotrypsinogen | 2 | 25 | 84 | 0.738 |
| Lysozyme | 2 | 13,6 | 104 | 0.9937 |

Fig.:Calibration Curve of Known Molecular Weight Proteins by Gel Filtration Chromatography.

Example(1):

Fig: Gel filtration profiles of alkaline RNase from Soft Tissue  on sephadex G -75 column.

Table: Isolation of Alkaline RNase Forms from Soft Tissue

| Groups | Vol. (ml) | Protein (mg/ml) | Activity (U/ml ) | Specific Activity (U/ mg) | Total Activity ( U ) | Yield ( % ) | Fold of Purification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Crude | 2 | 3.613 | 16.210 | 4.486 | 32.42 | 100 | 1 |
| Alkaline RNase Forms |  |  |  |  |  |  |  |
| 1 | 6.5 | 0.039 | 0.412 | 10.564 | 2.678 | 8.260 | 2.354 |
| 2 | 6.5 | 0.032 | 0.699 | 21.843 | 4.543 | 14.012 | 4.869 |
| 3 | 9.75 | 0.031 | 1.976 | 63.741 | 12.844 | 39.617 | 14.208 |
| 4 | 6.5 | 0.056 | 0.497 | 8.875 | 3.230 | 9.962 | 1.978 |
| 5 | 6.5 | 0.029 | 0.311 | 10.724 | 2.021 | 6.233 | 2.39 |

Activity (U/ml )

Specific Activity(U/mg)  =

Protein (mg/ml)

Total Activity( U ) = Activity (U/ml ) × Vol. (ml)

Total Activity of  partial purified

Yield( % ) =	 × 100

Total Activity of crude part

Specific Activity of  Partial Purified

Fold of Purification =

Specific Activity of Crude Part

Example(2):

Table:Purification Steps of Human Erythrocytes SOD from Control.

| Sample | Vol. ml | Total SOD (U) | Total Protein (mg) | Specific Activity (U/mg) | Yield % | Fold of Purification | Total Hb. (g) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Hemolysate | 195 | 9613.4 | 24031.2 | 0.4 | 100 | 1 | 7.64 |
| After ethanol-chloroform treatment | 200 | 9024.6 | 265.4 | 34.0 | 93.87 | 85 | - |
| After K2HPO4 | 70.5 | 8350.4 | 49.61 | 168.32 | 86.86 | 420.8 | - |
| After acetone | 8 | 8260.6 | 4.02 | 2054.8 | 85.92 | 5137 | - |
| After gel filtration chromatography | 15 | 8024.3 | 1.234 | 6502.67 | 83.469 | 16256.67 | - |

Fig.:Fraction Absorbance at 280 nm After Gel Filtration.

Crude                      After ethanol-                  After                        After              After gel filtration

Control                          chloroform                       K2HPO4