In order to minimize sample dilution use a maximum sample volume that gives the resolution required between the peaks of interest.

The ratio of sample volume to column volume influences resolution, where higher sample volume to column volume ratios give lower resolution.

Resolution in gel filtration increases as the square root of bed height. Doubling the bed height gives an increase in resolution equivalent to 2= 1.4 (40%). For high resolution fractionation long columns will give the best results and a bed height between 30–60 cm should be satisfactory. Sufficient bed height together with a low flow rate allows time for all 'intermediate' molecules to diffuse in and out of the matrix pores and give sufficient resolution.

Media Selection

Chromatography media for gel filtration are made from porous matrices chosen for their inertness and chemical and physical stability. The size of the pores within a particle and the

particle size distribution are carefully controlled to produce a variety of media with different selectivities. Today's gel filtration media cover a molecular weight range from 100 to 80 000 000, from peptides to very large proteins and protein complexes.

Superdex is the first choice for high resolution, short run times and high recovery.

Sephacryl is suitable for fast, high recovery separations at laboratory and industrial scale.

Superose offers a broad fractionation range, but is not suitable for large scale or industrialscale separations.

After deciding upon Superdex, Sephacryl or Superose, select the medium with the fractionation range that covers the molecular weight values of interest in your sample. In cases where two media have a similar fractionation range: select the medium with the steepest selectivity curve for best resolution of all components in the sample. When you are interested in a specific component, select the medium where the log molecular weight of the target component falls in the middle of the selectivity curve.

Sephadex is ideal for rapid group separations such as desalting and buffer exchange. Sephadex is used at laboratory and production scale, before, between or after other chromatography purification steps.

Sephadex G-25 is recommended for the majority of group separations involving globular proteins. This medium is excellent for removing salt and other small contaminants away

from molecules that are greater than Mr 5 000.

Sephadex G-10 is well suited for the separation of biomolecules such as peptides (Mr >700) from smaller molecules (Mr >100).

Sephadex G-50 is suitable for the separation of molecules Mr >30 000 from molecules Mr <1 500 such as labeled protein or DNA from free label.

Sample Preparation

Samples must be clear and free from particulate matter, particularly when working with bead sizes of 34 μm or less. The sample should be fully dissolved. Centrifuge or filter to remove particulate material.

Sample Clarification

Centrifugation and filtration are standard laboratory techniques for sample clarification and are used routinely when handling small samples.

It is highly recommended to centrifuge and filter any sample immediately before chromatographic purification.

1-Centrifugation:

Centrifugation removes lipids and particulate matter, such as cell debris. If the sample is still not clear after centrifugation, use filter paper or a 5 μm filter as a first step and one of the filters below as a second step filter.

•For small sample volumes or proteins that adsorb to filters,  centrifuge at 10 000 g for 15 minutes.

• For cell lysates, centrifuge at 40 000–50 000 g for 30 minutes.

•Serum samples can be filtered through glass wool after centrifugation to remove any remaining lipids.

2-Filtration:

Filtration removes particulate matter. Membrane filters that give the least amount of nonspecific binding of proteins are composed of cellulose acetate or PVDF.

For sample preparation before chromatography, select a filter pore size in relation to the bead size of the chromatographic medium.

Nominal pore size of filter       Particle size of chromatographic medium

1 μm                                                     90 μm and upwards

0.45 μm                                                         34 μm

0.22 μm                             3, 10, 15 μm or when extra clean samples or  sterile filtration is required

Sample Buffer Composition

The pH, ionic strength and composition of the sample buffer will not significantly affect resolution as long as these parameters do not alter the size or stability of the proteins to be separated and are not outside the stability range of the gel filtration medium. The sample does not have to be in exactly the same buffer as that used to equilibrate and run through the column. Sample is exchanged into the running buffer during the separation, an added benefit of gel filtration.

Samples should generally not exceed 70 mg/ml protein. Dilute viscous samples, but note sample volume

Remember that viscosity varies with temperature.

Buffer Composition and Preparation

The most important consideration is the effect of buffer composition on the shape or biological activity of the molecules of interest. For example, the pH and ionic strength of the buffer and the presence of denaturing agents or detergents can cause conformational changes, dissociation of proteins into subunits, dissociation of enzymes and cofactors, or dissociation of hormones and carrier proteins.Select a buffer and pH that are compatible with protein stability and activity and in which the product of interest should be collected. Use a buffer concentration that is sufficient to maintain buffering capacity and constant pH. Use up to 0.15 M NaCl to avoid non-specific ionic interactions with the matrix . Note that some proteins may precipitate in low ionic strength solutions. Use volatile buffers such as ammonium acetate, ammonium bicarbonate or ethylenediamine acetate if the separated product is to be lyophilized.

Use high quality water and chemicals. Solutions should be filtered through 0.45 μm or 0.22 μm filters. It is essential to degas buffers before any gel filtration separation as air bubbles can significantly affect performance. Buffers will be degassed if they are filtered under vacuum.

Column and Media Preparation:

Sample Application:

Elution and Flow Rates:

Controlling Flow Rate:

Checking Column Efficiency

Column performance should be checked at regular intervals by determining the theoretical plate number and peak symmetry. Prepacked columns are supplied with recommended values.

Typical values for column performance:

Superdex: Efficiency N>10 000, Peak symmetry As = 0.70–1.30

Sephacryl HR: Efficiency N>9 000, Peak symmetry As = 0.80–1.50

1.Equilibrate the packed column in distilled water at a linear flow of   60 cm/h.

2.Inject acetone (10 mg/ml in water) in a volume equivalent to 0.2% of the total packed column volume.

3.Monitor UV absorbance 280 nm from the time of injection until the acetone peak has eluted and the signal has returned to baseline.

4.Calculate column efficiency i.e. the number of theoretical plates (N):

N = 5.54 (Ve / W1/2)2 × 1000/L

where

Ve = peak elution (retention) volume

W1/2 = peak width at half peak height

L = bed height (mm)

Ve and W1/2 are in same units

Calculate the symmetry factor (As):

As = b/a

where

a = first half peak width at 10% peak height

b = second half peak width at 10% peak height