Powerful Visualization Technique Produces Time Averaged  Cross-Section Of Multiphase Domain In A non-invasive Measurements

Abbas Ali Mahmood Karwi

Technical Institute of Babylon

Abstract

The dual source computed tomography setup has been developing in the central and has potential for use as anon-invasive tool for determining the phases holdup distribution. The two sources used are Co60 (22mCi) and & Cs137 (235mCi) with energies for Cs137 equal to 0.662MeV and Co60  with two energies equal to 1.332MeV & 1.173 MeV . The object to be scanned can be exposed to both the γ-beams simultaneously. Each source is located inside Source collimator device (SCD), purpose of which is to collimate the beam to give it a fan shape. The count rate of un-attenuated photons will be measured by these detectors which will be used in image reconstruction algorithms. This DE-DSCT set-up is capable of carrying out 197 views 21 projections per view , spatially independent projection measurements during one scan of complete 360° rotation of the source around the multiphase system. Calculated results showed a substantial convergence with the real measurement values of dose rates especially in the common points.

الخلاصة

ان نظام التصوير المقطعى ثنائى المصدر تم تطويرة ليستخدم فى تصوير سلوك المائع فى حالتة الساكنة والمتحركة فى وسط ما. والمصدران المستخدمان هما كوبلت-60 ذو نشاط اشعاعى يساوى (22mCi) و سيزيوم-137 ذو نشاط اشعاعى يساوى (235mCi) مع طاقة سيزيوم-137 تساوي 0.662 ميكالكترون فولت  ولكوبلت-60 اثنين من الطاقات يساوي 1.332ميكاالكترون فولت  و 1.173 ميكاإلكترون فولت. يتعرض النموذج  اثناء الفحص لاشعة كاما من كلا المصدرين في آن واحد. ويقع كل مصدر داخل حافظة من الرصاص الغرض منها هو تجميع الاشعة لتطلق بشكل مركز ومخروطى. ان معدل عدد الفوتونات المنطلقة والمستقبلة من قبل مجسات التحسس لة ائر كبير فى وضوح وعتمة صورة الطور داخل النموذج. يتحرك حامل المجسات خلال دورة فحص واحدة اى 360 درجة حوالى 197 حركة وكل حركة تتكون من 21 جزء. أظهرت النتائج المحسوبة تقارب كبير مع القيم الحقيقية لقياس معدلات الجرعة عند نقاط الفحص المحتملة.

Introduction

The DSCT unit is a research machine designed to quantitatively determine the time averaged phase  holdup [11] distribution of the phases in a dynamic gas-liquid-solid- system (GLS) or other multiphase flow systems. It has been designed to use two sealed point gamma ray sources (Johansen et al ., 1996; Kak, 1998; Hu et al., 2005; Johansen, 2005). Currently, it is equipped with Co60 (~50 mCi) and Cs137 (~300 mCi). Each source is housed in a Source Collimator Device which is made of lead (for Cs137) and tungsten (for Co60). A fan beam arrangement of source-detectors is used for measuring the transmission of the gamma ray photons across the multiphase experimental setup (Kumar, 1994; Kumar et al., 1995; Kumar  & Dudukovic, 1997; Kumar et al., 1997). The fan beam consists of a longitudinal section of a cone. Each point gamma ray source is placed at the vertex of a cone. Its detector array is placed along the bottom section of its cone and the multiphase experiential setup is placed in the middle. The sources are positioned at the geometrical center of their SCDs to provide maximum shielding. The setup is designed so that the multiphase experimental setup placed at the center is simultaneously exposed to gamma photons from both sources. A detector array is located at the side opposite to each source in the respective fan beams. These arrays are capable of counting the un attenuated photons or even the scattered photons of the gamma ray that pass through the multiphase experimental setup . Typically a window of energy is set and the counts that fall in this window are recorded. Before a DSCT experiment is be performed, it is first ensured that the multiphase experimental setup is operating at the desired conditions. The SCDs are opened, to turn on the source. The DSCT experiment called a scan is then started and typically runs for about  8 hours. At the end, the SCDs are closed to ‘turn off’ the sources. During the scan, for a given source position , the detectors array is made to move in an angular manner. For each motion, the computer and data acquisition system collect the gamma ray counts data. The detector plate is moved 21 times; hence, the 15 detectors are oriented at 21 angular positions each with respect to the source. This way, for a given position of the source, 315 angular positions are covered along the arc of the fan beam. During a single scan, 197 source positions are covered for each of the sources around the multiphase experimental system. The DSCT motions are completely automated and controlled by the computer.

Materials and Methods

Gamma ray sources and the Source Collimator Devices

Figures (1-12) are for equipments used in Fulton Lab .The Source Collimator Devices (SCDs) were designed by Dr. Charles Alexander at Oak Ridge National Laboratory (ORNL) as a part of the DOE-anaerobic digester grant and manufactured by the Machine Shop facility at ORNL (Lrdere et al., 1996; Mudde et al., 2005, Nikitidis et al., 1999; Nooralahiyan & Hoyle; Osullivan&Benac, 2007). Lead was used for the Cs137 SCD, and tungsten was used for the Co60 SCD (Kak  et al., 1998). The gamma ray beam emerges from the collimator window, which is always shut with a window wedge when the source is not in use. The wedge is secured by a wedge pin (Schaffer, 1970; Ollinger, 1994; Rapaport et al.,1995; Rados et al., 2005), as shown in Figure 1.

Figure 1. Front view of the Source Collimator Device provided by the manufacturer                                              (Oak Ridge National Laboratory).

Figure 1. Front view of the Source Collimator Device provided by the manufacturer                                              (Oak Ridge National Laboratory).

Figure 2 . Front view of the bottom section of the Source Collimator Device when it is dismantled. The window wedge and location of the point source when the device is opened

Figure 2 . Front view of the bottom section of the Source Collimator Device when it is dismantled. The window wedge and location of the point source when the device is opened

Figure 3. Front view of the Source Collimator Device with the top, mid, and bottom sections assembled

Figure 3. Front view of the Source Collimator Device with the top, mid, and bottom sections assembled

Figure 4. Side view of the Source Collimator Device

Figure 4. Side view of the Source Collimator Device

The point source is located inside the lower half of an arming rod. In use, the arming rod is at the axial center of the SCD. Figure 2 shows the disassembled SDC. The window wedge is placed so that its apex occupies the axial center of the device, the location the source will be lowered to when the source is to be ‘turned on’ or ‘opened’. Figure 3 shows the SCD partially assembled and labels the window through which the gamma ray beam appears. This window is perfectly aligned with the detector lead collimators when the SCD is mounted on the DSCT setup. The source device has a wedge pin as a security device. To ‘open’ the source, the wedge pin is first removed. The window wedge is then removed with pliers. The top plate cover is removed by loosening the bolts attached to it. The removal of the top plate cover makes the plug retaining plate accessible. This plug retaining plate is attached to the arming rod that has the source in it. This externally threaded plate is lowered with a custom tool until the source is in the lower section of the SCD and centered in collimator window.

Shielding On DSCT

The following shielding components are used in the DSCT setup. 1. Source Collimator Device (lead or tungsten) (Dixon,1951; Sanchez & McConnel, 2000). The Source Collimator Devices (SCDs) are 6 inches in diameter. The closest distance from the outer surface of the SCD to the source is source is 3 in. for Cs137 and 3.5 in. for Co60. Hence, these are the minimum shielding thicknesses of the SCDs. Tungsten collimator for the Co60 has an additional lead shield of ½ in. thickness. 2-External Beam Collimator (lead): External Beam Collimators have been attached to each of the SCDs with the aid of annex tension from the upper section of the collimator strap, so that they are in line with the collimator window, as shown in Figure 4. These collimators, 2 inches in length at their center, make the gamma ray beam shorter in height as it emerges from the collimator window. 3. Extended Window Wedge: The window wedges supplied by the manufacturer are not used on the SCDs when the external source collimators are attached to them, because they are short and are inaccessible through the extended collimator. Extended Window Wedges made of lead, are used for both the SCDs while they are installed on the DSCT apparatus. 4. Detector Array Lead Shield: The Detector Array Lead Shields are the major shielding component of the DSCT setup as shown in Figure 7. These are curved lead blocks 2.5 inch thick and 3 inch tall that are placed in front of the detectors. They have fifteen  1 inch diameter holes in them to accommodate the Detector Lead Collimators.. Only photons that are incident on the aperture of the collimator pass through unhindered to the detectors.  5. Detector Lead Collimators: The Detector Lead Collimators have an aperture in the center. Two different sizes of apertures are used (small and large). The small aperture collimator is used by default. When a detector is not used in one of the slots in the detector array, a blind collimator (a collimator without an aperture) is inserted for that slot in the detector shield. 6. Extenders attached to Detector Array Lead Shield:  Lead extenders (shown in Figure 8), are placed at the edges of the detector array lead.  7. Lead Beam Stops: The Lead Beam Stop for each detector (indicated in Figures 9 and 10) is made of ½ in. thick lead and has an area of 1cm x1 cm.

Figure 5. View of the SCD showing the Set screw and the Source holder outer section. (Part 5)

Figure 5. View of the SCD showing the Set screw and the Source holder outer section. (Part 5)

Figure 6 . Top view of the SCD showing Part 12, the top plate

Figure 6 . Top view of the SCD showing Part 12, the top plate

Figure 7. Photograph of the DSCT setup without the Source Collimator Devices

Figure 7. Photograph of the DSCT setup without the Source Collimator Devices

Figure 8. Detector Array Lead Shield extenders, made of lead, are secured by bolts to the detector

array plate on either side of detector shields. They provide additional shielding at the sides

Figure 8. Detector Array Lead Shield extenders, made of lead, are secured by bolts to the detector

array plate on either side of detector shields. They provide additional shielding at the sides

Figure 9. Schematic of the fan beam arrangement of the Detector Lead Collimators with the Detector Array Lead Shield

Figure 9. Schematic of the fan beam arrangement of the Detector Lead Collimators with the Detector Array Lead Shield

Figure 10. Source Collimator Device and detector arrangement.The shaded triangular region in the middle shows a section of the gamma ray fan

Figure 10. Source Collimator Device and detector arrangement.The shaded triangular region in the middle shows a section of the gamma ray fan

Anticipated dose rates

1. The local dose rates in and around the DSCT setup when both SCDs are ‘turned on’. The highest possible dose received is at Location I, full body exposure with these dose rates at any location in the structure of the DSCT setup and around it. Since locations V and VI are beyond the aperture of the detector collimators as shown in Figure 11, the area of exposure for dose rates is same as the aperture of the collimator.

2- Occupational dose rates when the SCDs are ‘turned on’. The occupational dose has been calculated for a distance of 3 ft from the edges of the DSCT setup.

3- Public dose rates. These doses have been calculated assuming the concrete walls and the doors are absent, and the fan beams are directly oriented towards the wall. The wall thickness and an additional distance of 30 cm (11.81 in.) have been considered to determine the distance for these calculations as shown in Figure 12 . The shielding by the NaI crystal in the detector has been accounted . Since attenuation data for parts of the detector excluding the crystal was unavailable, attenuation by the crystal alone has been included. DSCT operation the public dose is calculated to be less than 2 mrem/hr. sight of the source when it is ‘‘turned on’’. These dose rates will be received, if one is .Since the dose rates calculated here doesn’t account for the shielding from the block walls, doors, and the multiphase experimental column in the center of the DSCT setup, the measured values will be less than 1mrem/hr in any one   hour . Dosimeters will be placed on the walls on the North, the West and the South sides, and on the mesh on the East side such that they are in line with the DSCT set up. These will record the dose received at the periphery of the room due to the DSCT experiments. These dose records will be a good estimate (after accounting for attenuation by the walls) of the public dose rates. These dosimeters will be returned every quarter to Radiation Safety, and the new ones issued will be placed at the same location.

4- Dose rates related to closed SCDs when the source is ‘turned off’. These doses will be received in the region behind the SCD even when the source is ‘turned off’

Figure 11. Schematic showing a top view of (DSCT) setup.

Figure 11. Schematic showing a top view of (DSCT) setup.

“RPT”  Setup

191 in

159 in

310 in

82 in

121 in

241 in

217 in

77 in

126 in

Wire Mesh

1

2

5

6

7

8

9

10

11

13

Cs-137

14

16

155

17

19

18

21

20

22

24

23

25

Electronics cabinet

“PBR” Setup

Black board

14

15

16

17

3 ft

3 ft

3 ft

3 ft

E

W

N

Wire mesh south door Locked

204 in

4

3

North Door

12

RPT setup

X

S

“RPT”  Setup

191 in

159 in

310 in

82 in

121 in

241 in

217 in

77 in

126 in

Wire Mesh

1

2

5

6

7

8

9

10

11

13

Cs-137

14

16

155

17

19

18

21

20

22

24

23

25

Electronics cabinet

“PBR” Setup

Black board

14

15

16

17

3 ft

3 ft

3 ft

3 ft

E

W

N

Wire mesh south door Locked

204 in

4

3

North Door

12

RPT setup

X

S

191 in

191 in

6

6

* Left side / (1ft) from Cs source –point (14) : * Behind / (1ft) from Cs source –point (15) : * Right side/ (1ft) from Cs source –point (16)  *Above / (1ft) from Cs source –point (17) * Left side / (1ft) from Co source –point (20) : * Behind / (1ft) from Co source –point (19) : * Right side/ (1ft) from  Co source –point (18)  *Above / (1ft) from Co source –point (21)

* Left side / (1ft) from Cs source –point (14) : * Behind / (1ft) from Cs source –point (15) : * Right side/ (1ft) from Cs source –point (16)  *Above / (1ft) from Cs source –point (17) * Left side / (1ft) from Co source –point (20) : * Behind / (1ft) from Co source –point (19) : * Right side/ (1ft) from  Co source –point (18)  *Above / (1ft) from Co source –point (21)

Figure 12. G-2 Fulton Hall room

Figure 12. G-2 Fulton Hall room

Operational Controls

3.1 Rope barrier around DSCT setup: The DSCT setup will have a rope barrier  around it at a minimum distance of 3 ft from the DSCT, or the 2 mrem/hr perimeter, which ever is greater. This is to demarcate the minimum distance the DSCT User needs to maintain from the DSCT setup when a DSCT experiment is on. further if the dose due to scatter is high near the DSCT setup.

3.2 Red Beacon: A red beacon is fitted on top of the DSCT’s computer tower. This is connected to a sensor box, placed on the circular source table (Figure 1). When either of the SCDs is opened to ‘turn the source on’ the sensor detects the radiation and the beacon is turned on. This sensor is set at threshold of 2 mrem/hr.

3.3 Postings: The following signs will be posted on all doors (North, South and West).                a- “Caution-Radioactive Material”.b-“Authorized Personal Only”.c-“Caution Radiation Area”.d-“Warning: Do not enter if the RED BEACON is on.”. e- “ This gate is to be shut at all times”-(For the Riser gate) On all sides of the DSCT setup’s frame. f-“ Caution High Radiation Area” - on all sides of the DSCT structure. g-“ Critical Shielding Component, not to be removed”.

3.4 Tools for Opening Source Collimator Devices (SCDs): The SCDs require a special key to open the top cover plate and a special screw driver to rotate the source plate and lower it down to the collimator window.

Procedure for Radiation Surveys

DSCT  radiation workers must conduct instrument surveys monthly and after any CT work, and must document these results properly. There must be no more than a 30 day gap between monthly surveys. If the lab contains no radioactive materials (RAM) or sealed sources, but the lab is still on active status as  defined by radiation safety, it is still required to perform and document monthly lab surveys. The instrument surveys are used to detect radiation and to identify areas of contamination. Radiation workers must routinely survey the place of work and themselves individually, both during and after their work. The areas to be surveyed are: areas of radioactive material use and storage, components of the DSCT setup, experimental setup and its components, and the equipment used with the DSCT instrument.

Instrument Surveys for Radiation Fields and Contamination

Radiation survey instruments [23 and 26] are available in areas where sealed sources/devices are used. In these laboratories, there are two different types of portable Geiger-Muller (G-M) surveymeters: one is a Bicron Model 50 with cylindrical G-M probe, and the other is a Ludlum Model 3 with a pancake G-M probe. Instrument surveys are valuable in identifying areas where radioactive contamination may be present and surveying the radiation field around the sealed sources/devices. In a typical survey, dose rate measurements are made in the vicinity of the sources/devices and throughout the laboratory. In addition, these portable survey instruments are used for a variety of general tasks:

1- To conduct routine area surveys of the laboratory. 2- To confirm the successful opening and shutting of the collimator containing the sealed source. 3- To survey shipments of sealed sources received. 4- To monitor hands, shoes, clothing, and the work area for contamination before leaving the area.

Portable Geiger-Muller (G-M) Survey Meters

G-M instruments are widely used for radiation survey work because of their reliability and low cost. The Bicron G-M probe is suitable for measuring exposure rates (mR/hr) in the vicinity of radioactive sources, but is relatively insensitive for detecting low levels of radioactivity associated with contamination. The Ludlum pancake G-M probe is much more sensitive to low levels of radioactive contamination, and the open side of the detector can be used for surface contamination measurements in counts per minute (cpm). The back side of the detector can be used like any G-M probe for exposure rate or dose rate measurements (mR/hr or mrem/hr). Figure 13 shows the Ludlum model3 survey meter with the pancake probe for general-purpose surveying.

Figure 13 . The Ludlum Model 3 Survey Meter with Model 44-9 G-M Pancake Probe

Figure 13 . The Ludlum Model 3 Survey Meter with Model 44-9 G-M Pancake Probe

Table 1. Dosimeter models

| Serial No | Dose equivalent | Error% | Calibration factor |
| --- | --- | --- | --- |
| E0740 | 41µSv | 0.0 | 1 |
| E0741 | 41µSv | 0.0 | 1 |
| E0742 | 41µSv | -2.4 | 1.03 |

Electronic pocket dosimeter (EPDTM ):

This instrument ideally suited for measuring the accumulated dose equivalent to X-ray of 20KeV  and used in medical service and other fields.The obtained data can be stored for as long as the battery is left in.There are three serial numbers for this device as shown in Table 1

Figure 14. Electronic pocket dosimeter instrument (EPDTM )

Figure 14. Electronic pocket dosimeter instrument (EPDTM )

Results and Discussion

Results

Table 2 presents the concentration in mR/h of 235mCi Cs137 and 22mCi C060 sources for twenty five samples examined  as locations by using ɤ-ray tomography. For locations 1-13 when the two sources are closed, the concentration is between  0.01-0.05 mR/h ,but when the Cs137 source is opened the concentration is between 0.02-0.05. Co60 source gives concentration between 0.02-0.07. When the two sources are in work the maximum concentration is at point 6 as shown in Figure 12. Table 3 presents the equivalent dose for locations 1ft from the two sources at points 14-21 which show the maximum rate of concentration at point 18 when the two sources in operation. Table 4 presents the equivalent rates of concentration for locations 22-25 for distance 3ft from the view plate with maximum rate at point 24 equal to 1.5 mR/h when the two sources are active.  Figures 15,16,17, and 18 shows the results obtained using the Ludlum Model 3 Survey Meter with Model 44-9 G-M Pancake Probe and  electronic pocket dosimeter instrument (EPDTM ) . Figures 19 and and 20  show equivalent dose rate results in (mR/h) and (rad/h) in 3D.

Table 2 . Equivalent dose rate for locations from 200 mCi Cs-137 & 50 mCi Co-60 sources /Distance from source to floor =41in

Table 2 . Equivalent dose rate for locations from 200 mCi Cs-137 & 50 mCi Co-60 sources /Distance from source to floor =41in

| Locations | Dose rate (mR/h)/Cs-137 closed | Dose rate (mR/h)/Cs-137 opened/Co-60 closed | Dose rate (mR/h)/Co-60 opened/Cs-137 closed | Dose rate (mR/h)/Co-60 +Cs-237opened |
| --- | --- | --- | --- | --- |
| 1 | 0.01 | 0.02 | 0.02 | 0.02 |
| 2 | 0.01 | 0.02 | 0.02 | 0.02 |
| 3 | 0.01 | 0.02 | 0.05 | 0.02 |
| 4 | 0.01 | 0.02 | 0.03 | 0.02 |
| 5 | 0.03 | 0.5 | 0.02 | 0.2 |
| 6 | 0.03 | 0.2 | 0.07 | 0.1 |
| 7 | 0.05 | 0.2 | 0.5 | 0.5 |
| 8 | 0.01 | 0.05 | 0.02 | 0.05 |
| 9 | 0.03 | 0.03 | 0.02 | 0.02 |
| 10 | 0.02 | 0.02 | 0.02 | 0.02 |
| 11 | 0.01 | 0.02 | 0.02 | 0.05 |
| 12 | 0.03 | 0.02 | 0.03 | 0.02 |
| 13 | 0.00 | 0.02 | 0.02 | 0.02 |

Table 3 .  Equivalent dose rate for locations (1ft) from 200 mCi Cs-137 & 50 mCi Co-60 sources/ Distance from source to floor =41in

Table 3 .  Equivalent dose rate for locations (1ft) from 200 mCi Cs-137 & 50 mCi Co-60 sources/ Distance from source to floor =41in

| Locations | Dose rate (mR/h)/Cs-137 closed | Dose rate (mR/h)/Cs-137 opened/Co-60 closed | Dose rate(mR/h)/Cs-137 opened/Co-60 closed (EPD)TM | Dose rate (mR/h)/Co-60 opened/Cs-137 closed | Dose rate (mR/h)/Co-60 +Cs-237opened |
| --- | --- | --- | --- | --- | --- |
| 14 | 0.05 | 0.8 | 1.0 | 0.1 | 0.07 |
| 15 | 0.05 | 0.5 | 0.5 | 0.2 | 0.06 |
| 16 | 0.05 | 0.5 | 0.5 | 0.1 | 0.08 |
| 17 | 0.05 | 1.0 | 1.0 | 0.1 | 0.08 |
| 18 | 0.02 | 0.4 | 0.5 | 0.4 | 1.5 |
| 19 | 0.01 | 0.4 | 0.5 | 0.7 | 1 |
| 20 | 0.02 | 0.3 | 0.5 | 0.4 | 1 |
| 21 | 0.02 | 0.4 | 0.5 | 0.3 | 1 |

Table  4.  Equivalent dose rate for locations (3ft) from 200 mCi Cs-137 & 50 mCi Co-60 sources/ Distance from source to floor =41in

Table  4.  Equivalent dose rate for locations (3ft) from 200 mCi Cs-137 & 50 mCi Co-60 sources/ Distance from source to floor =41in

| Locations | Dose rate (mR/h)/Cs-137 closed | Dose rate (mR/h)/Cs-137 opened/Co-60 closed | Dose rate (mR/h)/Co-60 opened/Cs-137 closed | Dose rate (mR/h)/Co-60 +Cs-237opened |
| --- | --- | --- | --- | --- |
| 22 | 0.05 | 0.5 | 0.1 | 0.2 |
| 23 | 0.02 | 0.5 | 0.3 | 0.2 |
| 24 | 0.02 | 0.5 | 0.4 | 1.5 |
| 25 | 0.07 | 1.0 | 0.4 | 0.1 |

Figure  15. Measured dose rate (mR/h) for G-2 Fulton Hall room when both sources are closed

Figure  15. Measured dose rate (mR/h) for G-2 Fulton Hall room when both sources are closed

Figure 16. Measured  dose rate (mR/h) for G-2 Fulton Hall room when Cs-137 open/Co-60 close

Figure 16. Measured  dose rate (mR/h) for G-2 Fulton Hall room when Cs-137 open/Co-60 close

Figure 17. Measured dose rate (mR/h) for G-2 Fulton Hall room when Cs-137 close/Co-60 open

Figure 17. Measured dose rate (mR/h) for G-2 Fulton Hall room when Cs-137 close/Co-60 open

1ft fft

2ft ft

Locations (1-13)

1ft fft

2ft ft

Locations (1-13)

Figure 18. Dose rate (mR/h) for G-2 Fulton Hall locations .Distance from source to floor =41in

Figure 18. Dose rate (mR/h) for G-2 Fulton Hall locations .Distance from source to floor =41in

Figure 19. Dose rate(mR/h) of G-2 Fulton Hall room locations

Figure 19. Dose rate(mR/h) of G-2 Fulton Hall room locations

Figure 20. Dose rate (rad/h) of G-2 Fulton Hall room locations

Figure 20. Dose rate (rad/h) of G-2 Fulton Hall room locations

Mass energy absorption/attenuation coefficient of water [1]

Mass energy absorption/attenuation coefficient of water [1]

Calculated results

Calculated results

Properties of water are assumed for calculating the dose absorbed by human body

Properties of water are assumed for calculating the dose absorbed by human body

Mass energy absorption/attenuation coefficient of lead [1]

Mass energy absorption/attenuation coefficient of lead [1]

ρlead=11.35

Mass energy absorption/attenuation coefficient of Tungsten [1]

Mass energy absorption/attenuation coefficient of Tungsten [1]

ρTungsten=19.35

Sources: Cs137 =235 mCi  and Co60 =22 mCi

Energy of gamma peak from Cs137 =0.662 Mev  ,( Branching factor=0.85)

Energy of gamma peak from Co60  =1.332 Mev  ,( Branching factor=1) & 1.173 Mev,( Branching factor=1)

Sources: Cs137 =235 mCi  and Co60 =22 mCi

Energy of gamma peak from Cs137 =0.662 Mev  ,( Branching factor=0.85)

Energy of gamma peak from Co60  =1.332 Mev  ,( Branching factor=1) & 1.173 Mev,( Branching factor=1)

ρNaI=3.67

Io[×[-----------(1)

Where; S:source strength(Ci) , E:Energy of gamma photon(Mev) , B.F:Branching factor of gamma photon, d=distance from source(cm)

ρs=surface density of substance receiving dose (

Io[×[-----------(1)

Where; S:source strength(Ci) , E:Energy of gamma photon(Mev) , B.F:Branching factor of gamma photon, d=distance from source(cm)

ρs=surface density of substance receiving dose (

Incident energy flux [2]

Incident energy flux [2]

Dose absorbed [2]

Dose absorbed [2]

Dose= Io [][1-exp

Where; Io: incident energy flux , :Absorption coefficient of water, ρwater:water density(

Dose= Io [][1-exp

Where; Io: incident energy flux , :Absorption coefficient of water, ρwater:water density(

Assuming 1 gm of water receives this dose in I cm2 × 1cm volume

,ing

Build up factor for Cs137 through lead [3- 5]

Build up factor for Cs137 through lead [3- 5]

B=A1×exp(-a1.µpb.T)+(1-A1)×exp(-a2.µpb.T)  --------------------------------------------------------------------------------------------(3)

A1=2.632 , a1=0.01465 , a2=0.136

B=A1×exp(-a1.µpb.T)+(1-A1)×exp(-a2.µpb.T)  --------------------------------------------------------------------------------------------(3)

A1=2.632 , a1=0.01465 , a2=0.136

Dose with shielding

Dose with shielding

Dose shield= Dose ×B × exp [ ---------------------------------------------------------------------(4)

:attenuation coefficient of lead ,ρpb:lead density ( , T: shielding thickness(cm)

Dose shield= Dose ×B × exp [ ---------------------------------------------------------------------(4)

:attenuation coefficient of lead ,ρpb:lead density ( , T: shielding thickness(cm)

Table 5. Input Data/Cs-137

Table 5. Input Data/Cs-137

| No | Variable | Symbol | value | Unit |
| --- | --- | --- | --- | --- |
| 1 | Cs-137 source activity | C1 | 235 | mCi |
| 2 | Distance -1 | d1 | 12.7cm=5in | Cm |
| 3 | Distance-2 | d2 | 33.02cm=13in | Cm |
| 4 | Distance-3 | d3 | 91.44cm=36in | Cm |
| 5 | Distance-4 | d4 | 71.2cm=28in | Cm |
| 6 | Distance-5 | d5 | 114.3cm=45in | Cm |
| 7 | Distance-6 | d6 | 120.65cm=47.5in | Cm |
| 8 | Distance-7 | d7 | 120.65cm=47.5in | Cm |
| 9 | Distance-8 | d8 | 189.23cm=74.5in | Cm |
| 10 | Distance-9 | d9 | 246.38cm=97in | Cm |
| 11 | Distance-10 | d10 | 513cm | Cm |
| 12 | Distance-11 | d11 | 520cm | Cm |
| 13 | Distance-12 | d12 | 185 | Cm |
| 14 | Distance-13 | d13 | 303 | Cm |
| 15 | Distance-14 | d14 | 7.62cm=3in | Cm |
| 16 | Distance-15 | d15 | 37.62cm=15.5in | Cm |
| 17 | Lead shielding distance-1 | TL1 | 6.35cm=2.5in | Cm |
| 18 | Lead shielding distance-2 | TL2 | 1.25cm=0.75in | Cm |
| 19 | Lead shielding distance-3 | TL3 | 7.62cm=3in | Cm |
| 20 | NaI Shielding distance | TNaI | 5.08cm=2in | cm |

Table 6. Output Data/Cs-137

Table 6. Output Data/Cs-137

| No | Location | Symbol | value | Unit |
| --- | --- | --- | --- | --- |
| 1 | LOCATION-I | X1 | 1.15 | rad/h |
| 2 | LOCATION-IA | X2 | 0.77 | - |
| 3 | LOCATION-IB | X3 | 0.59 | - |
| 4 | LOCATION-II | X4 | 3.6E-2 | - |
| 5 | LOCATION-III | X5 | 9.2E-2 | - |
| 6 | LOCATION-V | X6 | 1.02E-2 | - |
| 7 | LOCATION-IV | X7 | 1.02E-2 | - |
| 8 | LOCATION-VI | X8 | 2.55E-3 | - |
| 9 | OCCUPATIONAL | X9 | 4.3E-3 | - |
| 10 | PUBLIC- NORTH | X10 | 1.8E-3 | - |
| 11 | PUBLIC- SOUTH | X11 | 1.9E-3 | - |
| 12 | PUBLIC- EAST | X12 | 2.4E-3 | - |
| 13 | PUBLIC- WEST | X13 | 6.5E-3 | - |
| 14 | SHIELDING/IV | X14 | 9.8E-2 | - |
| 15 | SHIELDING/VI | X15 | 1.1E-2 | - |
| 16 | SHIELDING/OCCUPATIONAL | X16 | 7.6E-2 | - |
| 17 | SHIELDING/PUBLIC-NORTH | X17 | 7.6E-2 | - |
| 18 | SHIELDING/PUBLIC-SOUTH | X18 | 7.6E-2 | - |
| 19 | SHIELDING/PUBLIC-EAST | X19 | 7.6E-2 | - |
| 20 | SHIELDING/PUBLIC-WEST | X20 | 7.6E-2 | - |
| 21 | SHIELDING/BETWEEN SOURCE AND COLLIMATOR | X21 | 5.7E-2 | - |
| 22 | SHIELDING/BETWEEN SOURCE AND 30CM FROM COLLIMATOR SURFACE | X22 | 1.4E-2 | - |

Discussion

In estimating the radiation dose induced from Cs137 and Co60 sources, the maximum dose when 235mCi Cs137 is active is 1 mR/h at distance 1 ft for points 14,15,16,17,18,19,20,and 21 with values 0.8,0.5,0.5,1,0.4,0.4,0.3,and 0.4 ,while these values decreased at 3ft distance around view plate from  1 rad/h as maximum value to 0.5 rad/h as minimum value, for other locations 1-13 as presented in Table 2 the dose is between 0.02-0.03 which is not affected during the operation period .22mCi Co60 is lower in magnitude than the values of Cs137 .When the test is start and the two sources are opened the radiation dose reached maximum values especially near the view plate as shown in Figures 18, 19, and 20.So can say that the magnitude of dose increased maximum rates when the two sources are opened through the test especially in distance more the 3ft around view plate.

References

DixonW.R.,1951. Build-up factor for transmission of Co60 gamma-rays through concrete and lead. Physical Review, Letters to the editor :489-499.

Garrett C., Whyte G.N., 1954. Build-up measurement on Cobalt-60 gamma radiation in iron and   lead. Physical Review, 95(4):889-891.

Hu B, Stewart C, Hale CP, Lawrence CJ, Hall ARW, Zwiens H, Hewitt GF. 2005.Development of an X-ray    computed tomography (CT) system with sparse sources:application to three-phase pipe flow isualization Experiments in Fluids 39(4):667-678

Johansen GA. 2005. Nuclear tomography methods in industry. Nuclear Physics A 752:696-705.

Johansen GA, Froeystein T, Hjertaker BT, Olsen O. 1996. A dual sensor flow imaging tomographic system. Measurement Science & Technology 7(3):297-307.

Kak A, Slaney M. 1998. Principles of Computerized Tomographic Imaging. New York: IEEE Press. 83-86 p.

Kemoun A, Cheng Ong B, Gupta P, Al-Dahhan MH, Dudukovic MP. 2001. Gas holdup in bubble columns at elevated pressure via computed tomography. International  Journal of Multiphase Flow 27(5):929-46.

Kumar S. 1994. Computed Tomographic Measurements of Void Fraction and Modeling of the Flow in Bubble Columns [Ph.D Thesis]: Florida Atlantic University

Kumar S, Duduković MP. 1997. Computer Assisted Gamma and X-ray Tomography: Applications to Multiphase Flow Systems. In: J. Chaouki FL, M.P.Duduković, editor. Non-Inavsive Monitoring of Multiphase Flows.

Kumar SB, Moslemian D, Dudukovic MP. 1995. A [gamma]-ray tomographic scanner for imaging voidage distribution in two-phase flow systems. Flow Measurement and Instrumentation 6(1):61-73.

Kumar SB, Moslemian D, Dudukovic MP. 1997. Gas-holdup measurements in bubble columns using computed tomography. AIChE Journal 43(6):1414-1425. Carson R. 1984. EM Reconstruction Algorithims for Emission and Transmission Tomography. Journal of Computer Assisted Tomography 8(2):306-316.

Ledere MC, Hollander JM, Perlman I. 1966. Table of Isotopes. New York: john Wiley & Sons, Inc.

McGinnisR.,: http://www.radprocalculator.com/Files/ShieldingandBuildup.pdf.

Mudde RF, Bruneau PRP, vanderHagen THJJ. 2005. Time-Resolved gamma-Densitometry Imaging within Fluidized Beds. Ind. Eng. Chem. Res. 44(16):6181-6187.

National Institute of Standards and Technology (NIST)

Water: http://physics.nist.gov/PhysRefData/XrayMassCoef/ComTab/water.html

Tungsten: http://physics.nist.gov/PhysRefData/XrayMassCoef/ElemTab/z74.html

Lead: http://physics.nist.gov/PhysRefData/XrayMassCoef/ElemTab/z82.html

Nikitidis MS, Tüzün U, Spyrou NM. 1999. Determination of Phase Velocities in Multi-Phase Flows  Using Dual Photon Gamma-Ray tomography Chemical Engineering Communications 175(1):3 - 24.

Nooralahiyan AY, Hoyle BS. 1997. Three-component tomographic flow imaging using artificial neural network reconstruction. Chemical Engineering Science 52(13):2139-2148.

O'Sullivan JA, Benac J. 2007. Alternating Minimization Algorithms for Transmission Tomography. Medical Imaging, IEEE Transactions on 26(3):283-297.

Ollinger JM. 1994. Maximum-likelihood reconstruction of transmission images in emission computed tomography via the EM algorithm. Medical Imaging, IEEE Transactions on 13(1):89-101.

Rados N, Shaikh A, Al-Dahhan M. 2005. Phase Distribution in a High Pressure Slurry Bubble Column via a Single Source Computed Tomography The Canadian Journal of Chemical Engineering 83:104-112

Rapaport MS, Gayer A, Iszak E, Goresnic C, Baran A, Polak E. 1995. A dual-mode industrial CT. Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 352(3):652-658.

Roy S, Chen J, Kumar SB, Al-Dahhan MH, Dudukovic MP. 1997. Tomographic and Particle Tracking Studies in a Liquid-Solid Riser. Ind. Eng. Chem. Res. 36(11):4666-4669.

Roy S, Kemoun A, Al-Dahhan MH, Dudukovic MP, Skourlis TB, Dautzenberg FM. 2004. Countercurrent flow distribution in structured packing via computed tomography. Chemical Engineering and Processing 44(1):59-69.

Sarantites DG.2007. Radioactivity and Radiation Safety. Lecture notes Chemistry-437.Pg-46    http://www.chemistry.wustl.edu/~courses/spring_07/ch_437/Chem-437_rev4-25.pdf

Sanchez L.C., McConnel P.E.2000. Estimation of shielding thickness for a prototype department of energy national spent nuclear fuel program transport cask. Sandia Report (SAND 2000-1595). P-60. http://www.osti.gov/bridge/servlets/purl/759467-8TPizo/webviewable/759467.pdf.

Schaffer HE. 1970. Algorithms 369: Generation of random numbers satisfying the poission distribution Communication of the assocations for computing machinery 13(1).