MATLAB Assisted Analog ECG System Simulator

| Nasseer Moyasser Basheer | Omar Hatif Mohammed |
| --- | --- |
| Technical college of Mosul/ | Technical college of Mosul/ |
| Computer engineering Techniques department nasseer.58@gmail.com | Computer engineering Techniques department omerhatif@gmail.com |

Abstract

ECG machines play an important role in the primary diagnosis and survival analysis of heart diseases. Electrocardiography has had a profound influence on the practice of medicine. Large amount of ECG machines are the traditional analog ones, here simulation of an analog ECG is considered trying to get closer look at its design and aiming to better understanding and possible circuit modifications. This simulator is built using two major sections, the first one is used to make the electrodes signal (stander limb leads connection) suitable to be used by MATLAB, formed from an instrumentation amplifier and a data acquisition card. The second part is the Simscape MATLAB library performing the signal filtering and amplifying. The chosen ECG circuitry is got from an ECG training kit, but it can be modified as needed. This simulator shows same ECG signal as got by the kit and gives an easy means to modify the design and note the signal at different stages of the simulator.

Keywords: ECG, Simulator, Data Acquisition, ActiveFilters

الخلاصة

لعبت اجهزة اشارة تخطيط القلب دورا مهما في تشخيص و تحليل امراض القلب. فتخطيط القلب له دور كبير في مجال العناية الطبية.عدد كبير من اجهزة تخطيط القلب هي من النوع التقليدي التناظري, لذا هنا تمت محاكاة هذا النوع من الاجهزة للحصول على صورة اوضح للدوائر المستخدمة و بهدف فهم عملها و احتمالية تطوير محتوياتها. تم بناء هذا المحاكي باستخدام جزئين رئيسيين, الاول تم استخدامه لتهيئة اشارة  الاقطاب(في ربط الاطراف القياسي) لتكون مناسبة للاستخدام في الحاسبة من قبل برنامج ماتلاب, و يتكون من مكبرقياسات و جهاز اكتساب البيانات. و يتكون الجزء الثاني من مكتبة Simscape في ماتلاب التي تتم  فيها عمليات الترشيح و التكبير للإشارة. تم اختيار دوائر جهاز تخطيط القلب لتكون متشابهة مع دوائر جهاز تخطيط قلب مختبري و يمكن اجراء تعديلات عليها اذا اريد ذلك و معرفة تاثير ذلك على الاشارة. اشارة تخطيط القلب الناتجة هنا مشابهة تماما لاشارةجهاز التخطيط المختبري مع توفر وسيلة سهلة لاجراء تعديلات على كل الدوائر و كذلك امكانية معرفة شكل الاشارة عند المراحل المختلفة.

الكلمات الدالة : اشارة تخطيط القلب, محاكاة, اكتساب البيانات, مرشحات

- Introduction

ECG signal is one of the most important biomedical signals, this signal is a graphical representation of cardiac activity and it uses the primary measure for identifying various heart diseases and heart abnormalities [AntoniBayés de Luna, 2008]. The ECG signal in general contains: P wave, QRS complex and T wave. Each wave represents a unique activity of the heart as the following:

1. The P wave records atria contraction, the average duration of p wave is about 0.1 seconds, and its voltage is 0.1 to 0.3 mV [Joshi, 2006].

2. The QRS complex records ventricle contraction of duration 0.05 to 0.1 second [Joshi, 2006] it consists of three consecutive waves Q, R and S waves. Q wave is a small negative wave with voltage of average value of 0.06 to 0.19 mV. The R wave is the largest component in the ECG, its voltage is 0.5 to 1.8 mV. S wave is the next downward deflection of QRS its voltage is about 0.4 mV [Khurana, 2008].

3. The T wave records ventricle diastole, it lasts for about 0.27 second and has voltage of 0.15 to 0.5 mV [Joshi, 2006].

- Ecg Leads Placement

The ECG signal is recorded by placing an array of electrodes at specific locations on the body surface. Conventionally, electrodes are placed on each arm and leg, and six electrodes may also placed at defined locations on the chest. There are three basic types of ECG leads: (i) standard limb leads, (ii) augmented limb leads, and (iii) Chest leads. Twelve leads may be utilized from these three types, each lead views the heart at a unique angle, enhancing its sensitivity to a particular region of the heart. The more views we have, the more information we have, so when analyzing the ECG lead recording will differ in shape as well as in amplitude, certain features within ECG waveform are more clearly visible in one configuration than in another. The chest leads are more useful to detect abnormalities such as anterior or posterior myocardial infraction and differentiate left and right branch block, while the limb leads (Standard and augmented limb leads) enable to determine cardiac axes and allow to determine the presence of  an inferior myocardial infraction [Steve Bowbricket al., 2006].

  - Standard Limb Leads ( Bipolar Leads):

It's one of the most widely used methods to record ECG signal developed by Willem Einthoven in 1902[Kathryn A. Kalanick, 2011].In this type electrodes are applied to the left leg, the right and the left arms.The bipolar leads represent a potential difference between any two selected electrodes.Three ECG leads can be used, any ofthem may give an ECG signal (between two electrodes) with minor differences between them [C. Raja Rao et al., 2001],the leads are marked as lead I, lead II and lead III.

  - Augmented Unipolar Limb Leads:

The augmented unipolar limb leads produces 50% higher voltages than the standard limb leads, this type founded by Frank Wilson in 1934[Kathryn A. Kalanick, 2011].  Each Augmented lead is recorded between a single exploratory electrode and central terminal (indifference electrode), in this type, three ECG leads are used [C. Raja Rao et al., 2001], the leads here are referred asaVR, aVL and aVF.

  - Unipolar chest Leads (Precordial Leads):

The unipolar chest lead is developed by the American Heart Association in 1938[Kathryn A. Kalanick, 2011]. In this type the ECG signal is obtained from the two arms and the left leg as well as from six chest points generally has the following notations (V1, V2, V3, V4, V5 and V6).

- Building The Ecg System

ECG signals taken across the skin are very weak signals, with low signal to noise ratios. ECG systems are usually built to extract the ECG signal and to remove the noise. There are two types of ECG systems: Traditional (analog) ECG systems and Digital ECG systems.Analog ECG systems are built by using analog components while the digital ECG systems use digital components such digital filters. This paper describes how to simulate the analog filters of an ECG system inside the computer by using MATLAB program, this simulation will help students to learn how the ECG system works, how to troubleshoot it, and to study the ECG characteristics.

To build an ECG System, the following ECG signal characteristics must be considered:

- The amplitudes of ECG signals collected on the body surface have amplitudes less than 10 mV [Shakti Chatterjee et al., 2010].

- The frequency range of ECG signal is 0.1 Hz to 100 Hz, so to capture all the information of ECG it should be sampled at 200 Hz or more [Udayashankara, 2010].

- An ECG may be corrupted  by different noise sources :

- Power line interference (50/60 Hz): is caused by improper grounding of the ECG equipment and interference from nearby equipment Electrode noise contact [Adam Gacek et al., 2011].

- EMG Noise (20-10KHz): this noise comes from the electrical activity of skeletal muscles during periods of contraction causing electromyography noise [TomazJarm et al., 2007].

- Motion Artifacts (2-5 Hz): the electrode-skin impedance can produce artifacts with electrode motion [Adam Gacek et al., 2011].

- Respiration noise: this noise comes from human respiration and have effective amplitudes with frequency components usually below 0.8 Hz [TomazJarm et al., 2007].

- Electrosurgical noise (100 KHz-1MHz): this type of noise is produced by the instruments that apply electrocauterization and it can completely destroy the ECG signal [Adam Gacek et al., 2011].

- Instrumentation Noise: normally generated by electronic devices used in the signal processing, e.g., producing saturation on the input devices that prevent the ECG signal from being captured [Adam Gacek et al., 2011].

The ECG system is implemented here by using four stages, as illustrated in figure 1. As the noisy ECG signal is delivered by the probes at low levels (about 1mV), it is necessary to use a high precision data acquisition stage with a resolution of 20 bits and minimum sampling rate of 200samples per second,( as simulated using MATLAB to decide the project requirements ).Such a data acquisition card was difficult to obtain, so an instrumentation preamplifier stage is used before the available DAQ to eliminate the ECG signal and assure good ECG signal given to the MATLAB environment.

Fig 1. Block Diagram of the Proposed ECGSystem Simulator.

  - ECG Electrodes:

An ECG electrode is a device attached to the skin on certain parts of the patient’s body, generally the arms, legs, and chest during an electrocardiogram procedure, it detects electrical impulses produced every time the heart beats. The ECG electrode converts the potential differences due to ion flows inside the human body into a measurable voltage, there are three types of electrodes [Sun I. Kim et al., 2007]:

A. Adhesive Ag/AgCl Electrodes

B. Dry Electrodes

C. Capacitively coupled electrodes

A reusable clamp electrodes is used in this work, which are based on Ag/AgCl Electrodes[Claudio Becchetti et al., 2013].

  - Preamplifier (Instrumentation Amplifier AD620):

A Preamplifier is used in most ECG systems to amplify the probes voltage signals. But as the signals are noisy the noise will also be amplified. So it is a general procedure to use an instrumentation amplifier with a high CMNR (Common Mode Noise Rejection) to reject the similar noise components coming from the probes. Also according to the Lead I connection method the two arms signals are subtracted so the instrumentation amplifier is used as a difference amplifier to get the ECG signal. This will not remove all the noise because the probes does not carry exact noise components, the remaining noise is to be removed in the coming ECG filters. AD620 is used as the preamplifier in the ECG system as shown in figure 2. The AD620 is a low cost, high accuracy instrumentation amplifier that requires only one external resistor (RG) to set its gain from 1 to 10,000, the gain of AD620 is set to about 50X by letting RG equalsto 1 KΩ, figure 2 shows how to use the AD620 to get ECG signal, by using three electrodes for both arms and the right leg (as in standard limb leads connection).

Fig 2. AD620 as apreamplifier in ECG system with electrodes connected in the standard limb lead connection.

  - Data acquisition (DAQ):

Data acquisition is the process of sampling signals that measure real world physical activities and convert the resulting samples into digital numeric values that can be manipulated by a computer. DAQ is used to convert the analog ECG signal got from the preamplifier to a form that can be understood by the computer. PCI-6023E DAQ card from National Instruments is used, this data acquisition card has 16 single-ended or 8 differential channels, and its resolution is 12 bits with sampling rate 200 KS/s, it also has four input voltage ranges, and in accordance, this will change its precision as shown in table 1.

Table 1. Measurement Precision of the National Instruments PCI-6023E DAQ [National Instruments, 2005].

Figure 3 shows the ECG signal delivered at the front end of first Simulink block by the DAQ card, drawn by using MATLAB scope block.

Fig 3. The ECG signal as delivered to MATLAB environment.

  - ECG Simulink Filtering stages:

The ECG signal entered to the computer suffers from a lot of noise, from different sources, and with different frequency values. In order to display the ECG signal in a way that makes it useful for diagnosis purposes, these noise elements should be removed. The noise sources are those already given at the beginning of article 3.

Noise removal is achieved by using different types of filters, each has its cutoff frequency chosen to remove one of the components of the compound noise signal affecting the ECG main signal. These filter stages are manipulated using MATLAB SIMULINK (Simscape library). The circuits used in this section are built by reference to the biomedical measurement system set KL-73001[KL-73001, 2004].

    - High pass filter (HPF):

A HPF is used to remove thenoiselow frequencies that come from respiration and the baseline drift noise. In this work an active HPF (Sallen and Key) with cutoff frequency 1 Hz and order 2 is designed as in figure 4.

Fig 4. Sallen and Key High pass Filter (2ndOrder) [STEVE WINDER, 2002].

By letting C1 and C2 equal to 1 farad and using Butterworth table, an active HPF of cutoff frequency 1 Hz and order 2 can be designed by using these equations [STEVE WINDER, 2002]:

R1=ωnLP2QLP=σLP,R2=2ωnLPQLP=σLP2+ωLP2σLP

Where: R1&R2 are the values of filter resistors,ωnLP is the natural frequency,QLP  is the Q-factor, σLP is the real part of poleposition andωLPis the imaginary part of pole position.σLPandωLP are taken from Butterworth table [STEVE WINDER, 2002].

Then values of resistors and capacitors are denormalized, by using this equation [STEVE WINDER, 2002]:

Xc=12πFcC,R=R'2πFcC

Where: Xc is the Capacitive reactance,Fc is the cutoff frequency (1 Hz), C is the denormalized value of capacitor, R' is the normalized value of resistor and R is the denormalized value of resistor.

By letting all capacitors (C1 and C2) values equal to 100 nFthis results in resistors values as: R1= 1.1254 MΩ and R2= 2.2508 MΩ.

Figure 5 shows the ECG signal after removing the low frequency noise components (the respiration and the baseline drift noise).

Fig 5. ECG Signal output from HPF.

    - Low pass filter (LPF):

A LPF is used to remove the high frequencies noise that come from EMG (electromyography) and electrosurgical noise. In this work an active LPF (Sallen and Key) with cutoff frequency of 80 Hz and order 8 is designed by cascading four stages of Sallen and Key LPF as in figure 7. Figure 6 shows a single stage of Sallen and Key LPF:

Fig 6. Sallen and Key Low pass Filter (2ndOrder) [STEVE WINDER, 2002].

When resistors (R1 and R2) are equal to 1 Ohm, values of capacitors (in micro farads) can be obtained by using these equations [STEVE WINDER, 2002]:

C1=2Qωn=1σ,C2=12ωnQ=σσ2+ω2

Where: C1&C2 are the values of filter capacitors, Q is the Q-factor, ωn is the natural frequency, σ is the real part of pole position andωis the imaginary part of pole position.In the case of Butterworth filters,ωn=1and C2=σ, σ is taken from Butterworth table [STEVE WINDER, 2002].

The values of resistors and capacitors are denormalized, by using this equation [STEVE WINDER, 2002]:

C=C'2πFcR

Where: Fc is the cutoff frequency (80 Hz), R is the denormalized value of resistor, C' is the normalized value of capacitor and C is the denormalized value of capacitor.

If all resistors values equal to10 KΩ then the values of capacitors in micro farad will be: C1=0.2028, C2=0.1951, C3=0.2393, C4=0.1654, C5=0.3581 C6=0.1105, C7=1.0197 and C8=0.0388.

Fig 7. Low pass Filter (cutoff frequency 80 Hz, 8th order).

Figure 8 shows the ECG signal at the output of the LPF given above in figure 7.

Fig 8. ECG Signal output from LPF.

    - Gain amplifier:

This stage can be achieved by using a non-inverting amplifier with gain equals to (10), as seen in figure 9, it can be seen here that a non-inverting amplifier is used. Its input signal is free from many noise components that have been discarded by the previous HPF and LPF stages.

Fig 9. Non-inverting amplifier with a gain of 10.

Figure 10 shows the signal before and after the gain amplifier.

(a)

(b)

Fig 10. a) ECG signal before gain amplifier(with vertical axis multiplied by 4). b) ECG signal after gain amplifier.

    - Band reject filter:

In this part of the design, the band reject filter is applied to remove the 50 Hz power line noise, a Twin Tee Notch filter (band reject filter) of 50 Hz is designed as in figure 11:

Fig 11. Notch filer [STEVE WINDER, 2002].

By using equations below [STEVE WINDER, 2002], resistors and capacitors can be obtained to be as shown in figure 12:

R1=R2=12πFcC,R3=R12,k=1−14Q

Where: k is feedback factor,Q is the quality factor (equals 5) and Fc is cutoff frequency (50 Hz).

Fig 12. Notch filer of 50 Hz.

The final ECG signal obtained after the 50Hz notch filter is shown in figure (13) below:

(a)

(b)

Fig 13. a) Final ECG signal. b) Enlarging the final ECG signal with waves notations.

- Conclusion

Although this design was intended to be manipulated using MATLAB Simscape library only, but problem of non-availability of a suitable DAQ card, lead to using a hardware instrumentation amplifier. This simulator was suggested as a tool for understanding the operation and even learn how to maintain analog ECG systems. The filters and amplifier circuits may be tailored according to any specific ECG system following the previous goal.

The circuit used hare are standard normally used circuits for such application. Electrodes connected in standard limb leads connection. All circuits can be modified or designs changed to give better idea about its effect on ECG signal in each step. Using other lead connections (unipolar Limb and chest lead) shouldchange the ECGfront stage (instrumentation amplifier) only without changing the other stages.

The hardware preamplifier used before the DAQ is a low cost general purpose easily available instrumentation amplifier (AD620) so it will not degrade the versatility andbenefit of the whole simulator. Also this MATLAB library proves to give quite acceptable and useful results.

- References

Adam Gacek, WitoldPedrycz, (2011),"ECG Signal Processing, Classification and Interpretation: A Comprehensive Framework of Computational Intelligence", Springer

AntoniBayés de Luna, MiquelFiol-Sala, (2008),"Electrocardiography in Ischemic Heart Disease: Clinical and Imaging Correlations and Prognostic Implications", John Wiley & Sons.

C. Raja Rao,Sujoy K. Guha, (2001),"Principles of Medical Electronics and Biomedical Instrumentation", Universities Press

Kathryn A. Kalanick, (2011),"Phlebotomy Technician Specialist 2nd ed", Cengage Learning

KL-73001, (2004), “KL-700 Biomedical Measurement System”, http://www.eduware.com.pe/docs/otras_areas/KL-700.pdf

National Instruments, (2005), “NI 6023E/6024E/6025E FamilySpecifications”, http://www.ni.com/pdf/manuals/370719c.pdf

R A Joshi, (2006), "Questions & Answers In Physi"ology And Biochemistry ( Along With Mcq)", New Age International

Shakti Chatterjee and Aubert Miller, (2010),"Biomedical Instrumentation Systems", delmarcengage learning

Steve Bowbrick BSc and Alex Borg MD MRCP, (2006), "ECG Complete", Elsevier Health Sciences,

STEVE WINDER, (2002),"Analog and Digital Filter Design 2nd Edition", newness

Sun I. Kim, Tae Suk Suh, (2007), "World Congress on Medical Physics and Biomedical Engineering Volume 1", Springer

TomazJarm, Peter Kramar, AnzeZupanic, (2007),"11th Mediterranean Conference on Medical and Biological Engineering and Computing 2007", Springer

Udayashankara, (2010), "Real Time Digital Signal Processing", PHI Learning Pvt. Ltd