Reverse Injection Enthalpimetric Technique for the

Determination of Cerium(IV) and Sulpha Drugs

| Fadhil M.Najib College of Science | Ahmad M.Abdulla College of Science | Sardir Zewar Electronic Engineer. |
| --- | --- | --- |
| University of Sulaimani, - Sulaimani. Kurdistan Region, Iraq |  |  |

(NJC)

(Received on 7/8/2006)                                               (Accepted for publication on  9/5/2007)

Abstract

The new thermometric sensor developed by the present authors and appeared in (Talanta, 2007, 71(1),141-148) has been subjected to further applications.  In this work  Ce(IV) has been determined using Reverse Injection Enthalpimetry (RIE) , in which the analyte Ce(IV)  was added to excess reagent (sulphanilic acid).  For the determination of (4 - 35 μg.mL-1)  Ce(IV)  in the final  solution, a fixed volume between (0.1 – 0.8)mL. analyte  present in  . (0.5 M H2SO4), was injected into an excess of sulphanilic acid (20 ml. 1.25x10-2M)  in the thermometric cell.  Only one blank of  0.5 ml.(0.5 M H2SO4) was performed and its value was subtracted from all the results.  The  calibration curve  was linear with  R2 = 0.999 with a recovery  between 96 – 101% with  precision of about 2%, and t-test  showed no significant difference between the true value and the mean at 95% confidence  limit, indicating good validity of the method. The above procedure was reversed for the determination of sulphanilic acid and some sulpha drugs.  In this case a portion of 0.5 mL. of the analyte or the standard containing not less than 10mg of sulphanilic acid or the sulpha–drug is injected into an excess of Ce(IV)  as a reagent  in the thermometric cell and the procedure will be continued.

Keywords: Cerium (IV) and Sulpha Drugs Determination Thermometrically, Reverse Injection Enthalpimetry.

الخلاصة:

تم استخدام الكاشف الحراري الجديد المقدم من قبل نفس باحثي العمل الحالي والمنشور في(Talanta, 2007,71     (1),141-148) في مزيد من التطبيقات . في هذا البحث استخدم الكاشف لتقدير السيريوم الرباعي بطريقة الحقن العكسي لقياس الانثالثي وذلك باضافة Ce(IV)  المراد تقديره الى وفر من الكاشف سلفانيلك اسيد.لتقدير (4-35µg.mL-1 Ce(IV))  في المحلول النهائي، تم حقن حجم محدد من بين (0.1-0.8ml)  من المراد قياسه والمحضر في (0.5ml H2So4)  داخل وفر من سلفانيليك اسيد (20ml,1.5x10-2M) الموجود في الخليه المهيئة لقياس الانثالثي . كما وتم قياس بلانك واحد فقط مقداره (0.5ml of 0.5 M H2SO4)  وطرحت قيمته من جميع النتائج. كذلك وجد ان منحني المعايره كان خطاً مستقيماً (R2=0.9990) والاسترجاعية بمدى تراوح بين (96-101%) وبتطابقية بلغت قيمتها 2%  .

كما أظهر اختبار (t )  عدم جوهرية لفرق بين القيمة الحقيقية ومعدل القيم التجريبية عند حدود الثقة 95%  مما يدل على جدوى هذه الطريقة في تقدير السيريوم الرباعي . لقد استلمت معكوس الطريقة المذكورة اعلاه في تقدير السلفانيليك اسيد وبعض ادوية السلفا ، داخل وفر من (Ce(IV)) ككاشف داخل الخلية ، ثم اكملت العملية كما ذكر اعلاه.

Introduction

Under  conditions  of  invariant  heat  capacity,  the  change  in  temperature  in  an  ideal  adiabatic  system  represents  a  linear  measure  of  the  quantities  reacted  in  a  stoichiometric  process[1].  This  simple  principle  is  the  basis  of  (Direct  Injection  Enthalpimetry  “DIE”).  The measurement depends on the calibration between ( ∆T) and number of moles(i.e.the concentration) of the reacting  species [1,2].  The procedure and its success  depends  on  a  rapid  automatic  recording  of  the  “temperature  pulses”  following the injection of the  sample  into  the  reagent,  besides,  the  reaction  must  be  very  fast.

The  main  advantages  of  the  method  are;  that  it  is  not  necessary  to  standardize  the  reagent, provided it is added in sufficient excess to make the reaction  at least (99%)  complete[3],  and  the  speed  with  which  individual  determinations  may  be  carried  out.

(DIE)  technique  was  first  applied  for acid -base reaction , lead  and  magnesium  with  (EDTA)  as  interesting  examples  of  exothermic  and  endothermic  chelation  reactions [1].  Later it was  applied in  iodometry  for  the  determination  of  chlorine  in water[3] and  for  the  determination  of  olefins[4]  depending  on  measuring  the  amount  of  heat  liberated  from  the  hydrogenation  of  pure  olefin.  In all these methods and even until a few years ago[5] thermistors  were used as sensors.

In a previous paper [6], the authors have developed a new sensor for thermometric  methods, and applied for the determination of pKa values of difficult acids and for the kinetic thermometric  determination of Iron(III).

The present work  showing further applications on the importance  of this new sensor.  This time it is applied for the determination  of cerium(IV), sulphanilic acid and some sulpha drugs in a reverse way of the normal  DIE  method.

This modification [7], was named ( Reverse DIE),  by the present authors; here, instead  of  injecting  an  excess  of  a  reagent  to  the  analyte,  the  latter  will  now  be  injected  into  an  excess  of  the  reagent.

Although Ce(IV) has been determined by several good methods such as, inductively coupled  plasma  (ICP)[8 – 11], fluorescence and chemiluminescence [12-14],  X-ray, stripping  voltammetry[15,16] and spectrophotometric  methods [17-28], to the authors knowledge none of the thermometric methods has been used for its determination.  The previous work[6] showed  that  cerium(IV),  has  a  strong  effect  on  the  oxidation  of  sulphanilic.  This   phenomenon   was exploited  in  the  present  work in two ways.  In the first case excess of sulphanilic acid was used  for  the  determination  of  cerium(IV), while the procedure was reversed in second case, that was, by regarding cerium(IV) in excess, as a reagent for the determination of sulphanilic acid and some sulpha drugs.

Experimental:

Reagents:

All  the  chemicals  used  were  either  of  the  analytical  grade  or  the  highest  purity  available,  unless  otherwise  stated.

Usual  distilled  water,  protected  from  carbon  dioxide  was  used  for  all  preparations:

Sulphanilic  acid  solution  of  different  concentrations prepared  from  solid  substance ( 99% purity from Merck) as a standard,  in  an  acetate  buffer  (pH  =  4.75). Another type of this compound  (of  98% purity from Fluka was used as a synthetic sample.

Stock  solution  of  cerium  (1000  ppm)  was  prepared  from  solid  powder of cerium(IV) sulfate  tetra  hydrate   from  (Fluka ).  The  weighed  compound  was  dissolved  directly  in  (1:1)  sulfuric  acid  solution  completely,  then  the  solution  was  diluted  to  obtain  the  stock  solution  in  (0.5M  or  2M)  acid  solution.

Sulphacetamide eye drops (20%) from  Samara  Drug. Industrial  Company  (SDI),

Each  milliliter  of  the  original  drug  solution  contained (200 mg) sulphacetamide  sodium  salt. The  (10%) of this drug was obtained from  (Medico  Labs,  Syria).  Each  milliliter  of  the  original  drug  solution  contains  “100 mg”  of  sulphacetamide  and  used  as  a  sample.

Apparatus:

The home-made thermometric titration system, in the differential mode, including the construction of the new sensor has been described in a previous paper[6], a  micro syringe  (100 μL)  which  was  used  for  the  injection  of  small  amounts  of  the  reagent.

Recommended  Procedure:

a) Determination of Cerium(IV):

A (5ml)  portion  of  sulphanilic  acid  (0.25M)  which  was  considered  to  be  excess  was  placed  into each of the two thermometric  cells,  completed  with  distilled  water  to  (20ml).  The  cells  were  placed  in their  correct  positions,as shown in fig.(1) ( see  ref 6),  and  finally  covered  by  the  press-on  lid  in  such  away  that  the two identical  thermometric  sensors  have  dipped  in  the  solution.  The  vertical  stirrer  was  put  down  the  solution  and  sufficient  time  was  allowed  until  the  system  has  reached  thermal  equilibrium; which was watched  from  the  digital  avometer  reading.  The  ceric  sulphate  solution, standard or sample, between  (0.1 – 0.8ml)  was  then  injected  rapidly into  the sample-cell (but not to the reference-cell).  For  each  injection,  a  final  peak  high  was  taken  in  (mV)  from  the  chart  recorder  after  (1 min)  from  injection,  and  its  value  was  also  read  from  the  digital  avometer,  and  recorded  manually.

b) Determination of Sulphanilic Acid and Some Sulpha Drugs:

A (5ml)  portion (or excess) of  the  cerium(IV) solution was placed  into each of the two thermometric  cells,  completed  with  distilled  water  to  (20ml). The whole procedure was then followed as in (a) with one difference, that was, instead of injecting cerium(IV),  (0.5 mL.) of standards of sulphanilic acid or  samples, were now injected into the excess of cerium(IV) present in the cell.  For preparing  a calibration curve from a standard drug preparation , the solution was first diluted 1:1 then volumes between (0.1 – 0.7 mL.) were injected as mentioned before.

Results and Discussion:

a) Determination of Cerium(IV):

Effect of Acid Medium:

For  each  acid  type,  three  replicates  were  performed  with  injection  of  the  same  amount  of  ceric  sulphate  solution,  and  the  results  are  shown  in  table  (1). As  shown  in this table    sulphuric  acid was  the  best  medium  for  the  preparation  of  ceric  sulphate  solution.  This  was  due  to  its  quick  reaction,  which  was  one  of  the  important  condition  for  applying  (DIE)  technique,  in  addition  to  its  high  output  signal  compared  with  nitric  acid,  which  was  the  only  useful  acid  solution  after  sulphuric  acid.

The  acids,  HClO4,  H3PO4,  and  HCl,  were  not  suitable  because  of  lower  solubility  of  the  ceric  sulphate  in  these  acids  especially  at  low  temperatures  between  (10 – 15 oC)  which  was  the  optimum  condition  for  the  transistor  sensor  performance [6].  Ceric  sulphate  solution  in  hydrochloric  acid,  however,  was  unsuitable  owing  to  the  reduction  of  ceric  to  cerous[29],  with  simultaneous  liberation  of  chlorine:

Table (1):  Effect of the Acid Media of the Ceric Sulphate Solution on the Reaction Process.

| Acid solution of (7.5 x 10-3M) ceric sulphate | Sulphanilic acid | Output signal (mV) |
| --- | --- | --- |
| 2M H2SO4 2M HNO3 2M HClO4 2M H3PO4 0.5M HCl | 0.25M 0.25M 0.25M 0.25M 0.25M | 140 90 48 -- -- |

2Ce4+  +  2Cl–                 2Ce3+  +  Cl2

Phosphoric  acid  was  not  suitable  too  due  to  the  formation  of  a  sparingly  soluble  ceric phosphate precipitate (pKsp = 23.7) [30]. Therefore the best  medium was  that of sulphuric  acid  solution.

It  is  worth  mentioning  here  that  Ce(IV) has  almost  a  constant  reduction  potential  in  a wide  range  of  sulphuric  acid  concentration  between  (0.5 – 4M).

A  series  of  trials  was  conducted  for  different  concentrations  of  sulphuric  acid  between  (0.5 – 4M)   containing  a  fixed  amount  of  (7.5  x  10-3M)  ceric  sulphate,  according  to  the  procedure  given  earlier.  These  results,  indicated  that  as  the  concentration  of  sulphuric  acid  has  increased  larger  exothermic  signals  were  obtained.  It  is  well  known  that  dilution  of  sulphuric  acid  is  an  exothermic  process,  therefore  there  was  an  extra  heat  production  inside  the  thermometric  cell  in addition  to  the  heat  of  combination  of  ceric(IV) and  sulphanilic  acid.  This  effect  was   studied by the  injection  of  the  same  volume  of  the  analyte  Ce(IV) in  sulphuric  acid medium,  and  sulphuric  acid  alone without Ce(IV)  as  a  blank    into  the  cell  containing  sulphanilic  acid.  It  was  concluded  that  the  blank  was  more  effective  at  high  concentrations  of  the  acid. Accordingly  the  concentration  of  sulphuric  acid  as  low  as  (0.5M)  was  chosen,  although  higher  concentrations  have  still  given  slightly  better  sensitivity,  but lower  concentration  than  this will  negatively  affect the oxidizing power  of  Ce(IV).  At  high  acid  concentration  the  blank  output  signal  will  become  so  large  that  it  may  be  higher  than  that  of  the  analyte.

The  performance  of  a  fixed  concentration  of  (0.091M)  sulphanilic  acid  was also examined  in  different  acids.

The  experiments  involved  the  use  of  either  similar  or  different  acidic  media  for  both  sulphanilic  acid  in  the  thermometric  cell  and  the  injected  cerium  solution  and  detail  observations  were  given  in  table  2.

It  was  hoped  that  the  similar  media  will  eliminate  the  extraneous  heat  of  mixing  and  dilution,  shown  in  the  previous  section,  but  the  results  of  this  table  showed  that,  non  of  these  were  satisfactory.

No  obvious  explanations  could  be  deduced  for  this  behaviour,  except  the  solubility  factor.  Apart  from  HNO3,  sulphanilic  acid  was  not  soluble  in  all  the  acids  tested.  Unfortunately,  this  media  gave  the  lowest  output  signal  (42 mV)  compared  with  those  obtained  when  acetate  buffer  was  a  medium  for  sulphanilic  acid.  In  this  medium  Ce(IV) in  sulphuric  acid  gave  the  highest  signal  of (140 mV)  compared  to  (90 mV)  and  (48 mV)  in  cases  of  HNO3  and  HClO4  respectively.  Once  again  the  same  media chosen  before,  for  both  sulphanilic  acid  and  Ce(IV),  which  were  acetate  buffer  and  (0.5M  H2SO4)  respectively,  have  proved  to  be  optimum.

Blank  Optimization:

In  the  previous  experiments  of  this  part,  the  blank  signal  was  a  problem  and  has  been  subtracted  from  the  signal  of  the  analyte  for  each  run,  to  eliminate  the  heat  of  dilution.  This  was  because  of  injecting  different  volumes  of  ceric  sulphate  –  sulphuric  acid  mixture,  when  variation  in  Ce(IV)   ion  concentration  was  needed,  for  instance,  in  a  calibration  by  this  method.  The  problem  was  solved  as  follows:

Ceric  sulphate  solution  was  prepared  with  different  concentrations  in  the  (0.5M)  sulphuric  acid  solution, thus  for  each  subsequent  experiments,  the  volume  of  ceric  sulphate  solution  injected  will  be  kept  constant  at  (0.5ml).  In  this  way  only  one  blank  injection  will  be  performed,  by  injecting  (0.5ml)  of  (0.5M  H2SO4)  solution.  Hence  minimizing  the  steps  of  analysis  and  also  the small  volume  injected,  which  was  highly  desirable  condition  for  (DIE)  technique  to  minimize  or  eliminate  the  effect  of  heat  of  dilution.  The  calibration  curve  of  this  type  and  the  method  of  its  preparation  is given in in  the  following  section,  including  its  comparison  with  that  given  before  blank  optimization.

Table 2:  Effect  of  the  Type  of  Acid  Media  of  Sulphanilic  Acid  on  its  Oxidation

by  Ceric(IV)  Ion

| Media [a] 0.091M sulphanilic acid in the cell | Media [b] Ce(IV) solution injected | Outpt signal (mV) | Notes |
| --- | --- | --- | --- |
| 1. H2SO4 | H2SO4 | -- | a. There is reformation of crystals of sulphanilic acid in cold solution (10oC). b. Soluble. |
| 2. HClO4 | HClO4 | -- | a. Insoluble especially in cold solution. b. Soluble. |
| 3. HCl | HCl | -- | a. Insoluble. b. Soluble (Decolourization of yellow colour of cerium solution) |
| 4. HNO3 | HNO3 | 42 | a. Soluble. b. Soluble (Colour of complex: deeper). |
| 5. H3PO4 | H3PO4 | -- | a. Insoluble. b. Insoluble. |
| 6. Acetate buffer | Acetate buffer | -- | a. Soluble. b. Soluble (Decolourization of yellow colour of cerium solution). |
| 7. Acetate buffer | 2M HClO4 | 48 | a. Soluble. b. Soluble (Colour of complex: deeper). |
| 8. Acetate buffer | 2M HNO3 | 90 | a. Soluble. b. Soluble (Colour of complex: deeper). |
| 9. Acetate buffer | 0.5M HCl | -- | a. Soluble. b. Soluble (Decolourization of yellow colour of cerium solution). |
| 10. Acetate buffer | 0.5M – 2M H2SO4 | 140 | a. Soluble. b. Soluble (Colour of complex: intense pink). |

Preparation of the Calibration Curves and a Comparison Between them:

Two  types  of  calibration  curves  were  prepared.  The  first  one  was  performed  with  no  blank  optimization,  i.e.  each  concentration  had  its  own  blank.  In  the  second  calibration,  however,  only  one  blank  was  used  according  to  the  discussion  given  in  the  previous  section,  the  two  calibration  method,  are  given  below:

1)  Injection of different volumes  between(0.1 – 0.8 ml) of  a fixed  concentrate of ceric  sulphate  of  (7.5  x  10-3M)  in  the  final  (0.5M  H2SO4)  to a constant  amount  of  sulphanilic  acid  (20  ml of  6.25 x 10-2M).  For  each  injection,  the  blank  was  the  same  volume as that of ceric sulphate solution of  (0.5M)    H2SO4  solution,  and  the  resulting  signal  was  subtracted  from  it.  The  results  were  shown  in  (fig.2a)  for  the  concentrations  of  cerium  in  the  range  between  (2.5 – 21 μg.mL-1) in the final volume.

2)  In  this  procedure,  the  blank  optimization  was  applied. A series  of  different  concentrations  of  ceric  sulphate  between  (4 – 35 μg.mL-1 in the final volume)  present  in  a  fixed  volume  of  (0.5 ml  of  0.5M  H2SO4)  were  injected  into  (20ml  of  1.25  x  10-2M)  sulphanilic  acid  in  the thermometric  cell.  Only  one  blank  injection  of  (0.5ml)  of  the  (0.5M)  sulphuric  acid  solution  was  performed  in  duplicate  and  its average  value,  was  subtracted  from  all  the  results,  which  were  shown  in (fig.2b).

a

b

µ g Ce(IV)

ΔT C °

Fig (2) Calibration Curve :

a – Type 1 by successive injection of Ceric

Sulphate solution.

b - Type 2 by injection constant volume (0.5 ml)

of Ceric Sulphate solution

As  shown  in  figure  1,  both  types  were  analytically  acceptable,  but  type  (b)  was  more  convenient,  because  only  one  blank  was  performed,  with  a  better  linearity

(R2  =  0.999)  compared  to  type  (a)  (R2  =  0.994).

Slope  Measurement  Method  for  Preparation  of  a  Calibration  Curve:

A  series  of  injections  were  performed  according  to  the  procedure,  but  for  each  injection  the  (temperature – time)  curve  was  taken  for  final  measurement,  as  this  method  was  performed  in Catalytic Thermometric Titration (CTT)  technique [6].  The  slope  (degree oC sec-1) was  measured  for  each  injection  and  a  new  calibration  curve  was  prepared  by  plotting  the  slopes  vs.  (μg cerium)  injected..  The  comparison  of this  calibration  of  (R2 = 0.965) with  that drawn  from  peak  height  measurements  and  given  in  fig.(1b)  which  gave  a  better  linearity  of  (R2 = 0.999)  makes  one  preferring  the  latter.  This  was  probably  due  to  the  fast  reaction  between  cerium(IV)  and  sulphanilic  acid,  since  the  slope  measurement  represents  the  rate  of  reaction  and  is  more  applicable,  apparently,  to  medium  and  slow  reactions.  Therefore,  the  method  of  peak  height  measurement  was  more  acceptable  for  this  work  and,  perhaps  other  similar  fast reactions.

Effect  of  Interferences:

Interference  studies  on  the  determination  of  (40 μg)  cerium  showed  a  good  degree  of  selectivity.  The  interfering  ions  were  examined  individually  by  the  addition  of  concentrated  solutions  of  their  salts,  shown  in  table  3,  to  the  thermometric  cell,  containing  sulphanilic  acid  solution  as  usual.  The  addition  of  the  interfering  ion  was  continued  until  the  tolerance  level  was  reached.  In  some  cases  the  solid  salt  was  added  directly  to  the  sulphanilic  acid  solution.  The  tolerance  level  was  defined  as  the  maximum  ratio  of  the  interfering  ion  to

ceric(IV)   ion  concentration  producing  an  error  not  more  than  (5%).  Table 3 shows  the  tolerance  levels  for  the  cations  and  anions  examined.

Table 3:  The  Tolerance  Level  of  the  Interfering  Ions

| Ion | Added from | Tolerance, limiting ratio [Ion]/[CeIV] |
| --- | --- | --- |
| I– Sn(II) Fe(II) Cr(VI) Cr(III) Co(II) Mn(II) MnO4– Pb(II) Ba(II)* Ca(II) Ni(II) Cu(II) Zn(II) Hg(II) Hg(I) Fe(III) | KI SnCl2 (NH4)2Fe(SO4)2.6H2O K2Cr2O7 CrCl3.6H2O Co(NO3)2.6H2O Mn(CH3COO)2.4H2O KMnO4 Pb(NO3)2 Ba(NO3)2 Ca(NO3)2 NiCl2.6H2O Cu(NO3)2.3H2O Zn(NO3)2 HgSO4 in 1M H2SO4 Hg(NO3)2.2H2O FeCl3 | 10 10 250 250 250 250 250 250 800 1000 1000 1000 1000 1000 2000 5000 10000 |

*Some turbidity appeared in the solution.

The  table  showed  that  the  most  significant  interference  was  that  for  (I–  and  Sn(II).  An  attempt  was  made  to  eliminate  them  by  their  oxidation  with  (H2O2).  This  was  done  by  the  addition  of  (2ml)  of  (10%  H2O2)  to  the  thermometric  cell,  then  stirred  for  (2  min)  more  time  than  usual.  This  treatment  was  sufficient  for  increasing  the  tolerance  level  for  both  ions  to  (100).

Analytical  Application  of  the  Present  Method:

In  order  to  assess  the  analytical  application  of  this  work,  and  because  the  cerium  ore  was  rare,  and  not  obtainable,  a  synthetic  sample  was  prepared  by  mixing  (51.29 μg)  of  cerium(IV)   with  an  accumulation  of  the  interfering  ions  listed  in  table 3. For this purpose one  milliliter  from  each  of  the  interfering  ions  was  placed  in  the  thermometric  cell,  with  the  sulphanilic  acid  solution,  this made the final  solution    turbid  before  the  injection  of  ceric(IV)   ion.  The  ratios  of  the  concentration  of  each  interfering  ion  to  that  of  ceric(IV),  in  the  synthetic  sample  solution  were  indicated  in  table  4.

Table  4:  Ratio  of  the  Accumulation  Ions  to  Ceric(IV)   Ion  in  the  Synthetic

Sample  Solution  of  Cerium

| Ions | Ratio [ion]/[CeIV] |
| --- | --- |
| CoII, CrIII, FeIII, FeII CuII, MnII, SnII NiII, ZnII, CrVI, I– CH3COO– BaII, HgII, HgI, PbII, Cl–, SO42– NO3– | 2.5 5 10 20 100 |

The  above  experiment  was  also  repeated  without  the  presence  of  foreign  ions.  The  results  of  the  two  sets  of  these  experiments  were  statistically  evaluated  to  establish  whether  or  not  these  interferences  have  introduced  any  systematic  error  to  the  system,  and  were  shown  in  table  5.

Statistical  evaluation  for  table  (5)  by  testing  the  difference  between  the  mean  ()  and  the  true  value  (μ) showed a significant difference  at  95%  C.L.  (DOF = 6),  due  to  the  presence  of  the  accumulation  of  foreign  ions.  But  the  effect  was  still  acceptable  at  such  low concentration  of  Ce(IV).  The  mean  error  value  with  all  these  interferences

Table 5: Determination of Cerium in a Synthetic Sample, With and Without Foreign Ions

| No. of replicate | μg Ce injected | With foreign ions |  |  | Without foreign ions |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | μg found | Recovery% | E% | μg found | Recovery% | E% |
| 1 2 3 4 5 6 7 | 51.29 51.29 51.29 51.29 51.29 51.29 51.29 | 47.656 47.138 48.692 47.656 48.174 46.620 47.138 | 92.91 91.90 94.93 92.91 93.92 90.89 91.90 | – 7.09 – 8.10 – 5.07 – 7.09 – 6.08 – 9.11 – 8.10 | 52.318 52.836 50.246 50.764 52.318 51.800 50.764 | 102.00 103.01 97.96 98.97 102.00 100.99 98.97 | + 2.00 + 3.01 – 2.04 – 1.03 + 2.00 + 0.99 – 1.03 |
|  |  | = 47.58 |  | Average – 7.23 | = 51.55 |  | Average 0.56 |
| S.D = 0.697 , RSD = 1.46 |  |  |  | S.D = 0.986 , RSD = 1.91 |  |  |  |

was  (– 7.23%)  with  a  reasonable  precision  of  (RSD = 1.46).  When the foreign ions are not present, however, the  difference  is  not  significant,  at  (95%)  confidence  limit (DOF = 6),  showing  that  no  systematic  errors  was  detected  in  the  results.  The  table  showed  that,  a  good  degree  of  accuracy  and  precision  were  obtained,  mean  of  (E% = 0.56)  and  (RSD = 1.91),  compared  with  other  methods  for  the  determination  of  cerium indicated  before.

The  present  method  is  simple,  rapid,  adequately  sensitive,  and  moderately  selective,  in  addition  to  other  common  advantages  of  thermometric  methods.

A  Recovery  Test:

The  accuracy  of  the  method  was further studied by  performing multiple recovery  test,  that  was  by  calculating  recovery  of  the  original  quantity  of  cerium,  then  for  the  quantity  added  and  finally  for  the  total  (original + added)  as  follows:

1.  A  series  of  experiments  were  carried  out  by  injecting  (2.11 mg)  of  cerium  to  the  excess  sulphanilic  acid  in  the  thermometric  cell,  according to the  procedure. The  percentage  recovery  of  the  original  quantity  of  this  set  of  five  injections  was  found  in  the  usual  way.     2.  For  each  of  the  five  original  solutions  mentioned  in  (1)  above  further  (1.055  mg)  of cerium(IV)  was  also  injected  and  the  results  were  again  measured  from  the  peak  heights.  These  results  together  with  the  total  percentage  recoveries  were  presented  in  table 6.

The  recoveries  ranged  between  (96 – 101)%  for  the  original,  (99 – 101)%  for  the  added  quantity,  while  for  the  total  recoveries  was  between  (97 – 101)%.These  results  were  generally  showing  reasonable  accuracies  and  the  (t-test)  indicated  no  significant  differences  between  the  mean  ()  and  the  true  value  in  each  case  at  (95%)  confidence  limit, (DOF = 4). This  was  an  indication  for  the  presence  of  no,  or  very  small  systematic  error  in  the  results.  The  appearance  of  (+  and  –)  in  the  results  also  supporting that

Table  6:  Multiple  Recovery  Test

| No. of rep. | Original quantity of Ce(IV) |  |  |  | Recovery of the added quantity |  |  |  |  |  |  |  | Total recovery |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | mg Ce injd. | mgCe found | R% | E% | mg Ce injd. |  | mg Ce found |  | R% |  | E% |  | Total mg Ce found | R% | E% |  |
| 1 2 3 4 5 | 2.11 2.11 2.11 2.11 2.11 | 2.02 2.04 2.09 2.14 2.02 | 95.73 96.68 99.05 101.42 95.73 | – 4.27 – 3.32 – 0.95 + 1.42 – 4.27 |  | 1.055 1.055 1.055 1.055 1.055 |  | 1.041 1.044 1.051 1.068 1.062 |  | 98.67 98.96 99.62 101.23 100.66 |  | – 1.33 – 1.04 – 0.38 + 1.23 + 0.66 | 3.061 3.084 3.141 3.208 3.082 | 96.71 97.44 99.24 101.36 97.38 | – 3.29 – 2.56 – 0.76 + 1.36 – 2.62 |  |
|  |  | Mean 2.06 | S=0.05 RSD= 2.5 | Aver. - 4.27 |  |  |  | Mean 1.053 |  | S = 0.02 RSD=1% |  | Av. -0.17 | Mean 3.115 | S=0.06 RSD = 2% | Av. -1.57 |  |

b) Determination of Sulphanilic Acid and Some Sulpha Drugs:

The calibration curve performed between  ΔT and mg quantities of Sulphanilic Acid standard solution injected, is shown in fig. 3, with a high degree of linearity ( R2 = 0.9998).  This calibration was used for the determination of the purity of a sample of sulphanilic acid.

Another calibration curve was prepared from Sulphacetamide (20%) eye drops which gave

( R2 = 0.9995) and used for the determination of this drug of other makes and the results are shown in table 7.

#### Table (7):  Results  of  the  Determination  of  Sulphacetamide  10% Drug

#### and   Sulphanilic Acid by  the  Present  Method

| #### No. of repli. | #### mg sulphacetamide drug injected | #### mg Sulphacetamide found | #### R% Average | #### E % Average | #### Statistical evaluation |
| --- | --- | --- | --- | --- | --- |
| 1 | 50 mg | 50.375 | 100.6 | 0.60 | Test for the difference between the mean and true value showed that the difference is not significant at (95%) confidence limit (DOF = 5). There is no systematic error. |
| 2 | mg sulphanilic acid injected in 0.5 mL | mg sulphanilic acid found |  |  | Test for the difference between the mean and true value showed that the difference is not significant at (95%) confidence limit (DOF = 5). There is no systematic error S.D.=0.34 RSD=1.56 |
|  | 21.65 | 21.81 | 100.7 | 0.7 |  |

Table 7 shows that once again the new sensor is well suited for the determination of sulphanilic acid, sulpha druge and perhaps many similar samples with good accuracy and precision.

Acknowledgement:  The Authors would like to thank the Sulaimani University in Kurdistan

Region of Iraq, for supporting this work.

References:

1- Wasilewski, J.  C.,  Pel, P.T. S. and  Joseph  Jordan,  Anal. Chem.  1964, 36,2131.

2- Bark L.S. and  Bark, S.M, Thermometric  Titrimetry,  from  international  Series  of  Monographs  in  Analytical  Chemistry , Pergamon  Press Ltd., 1st Edition, 1969 ,68,  67883).

3-   Korda, .  P.  M.and Erdey,L.,  Talanta,   1970, 17, 1215– 1218.

4-Rogers, D.  W. ,  Anal.  Chem.  1971,43,1468  –  1471.

5- Mario,J.D.C., Mario,A.F.J. and Oswaldo,E.S.G.,Jr.,Brazil.Chem.Soc.,2002,13(5),692 – 694.

6-   Najib,F.M., Zewar,S and Abdulla,A.M ,   Talanta, 2007, 71(1), 141 – 148.

7-   Abdulla,A.M, A New Thermometric Sensor and its Application in Different

Enthalpimetric Techniques,2003,  a Ph.D Thesis, Sulaimani University.

8-   Takahashi, Y., Shimizu,H., Kogi, H. and Yoshida,H., Usui,A.  and  Nomura,M.,

Goldschmith,  2000,  Ir.  Of  Conference  Abstracts,2000, Volume  5(2),  981.

9-   Liang  –  Ping,D.,  Fenxi  Shiyanshi,  2000,  19,  67;  Chem.  Abstr. 2000 ,133,  217008h.

10- Qunzhou,B.,  Subiao,P.,  Bailim,H.,  Zhiguang,Z., and  Weiliang,G.,  Yansu  Kexue, 1999, 6,  64;  Chem.  Abstr. 1999, 131,  82253v

11- Shuai,Q., Bin,H., Qin,Y., Ruth,W. and  Jiang,Z.,   Ir., Anal.At. Spectrochem., 2000,15, 1413.

12- Jie N.,  Si,Z., Yang,J., Haung, X., and Yang,D.,  Mikrochimica  Acta,1997, 126.

13- Cui H,Li S,Lin X, Analyst 2001, 126(5),553.

14- Karim MM,Lee SH, Kim YS, Bae HS, Hong SB, J.Fluoresc.,2006, 10(1), 17.

15- Beng,K.S., and  Jing,Z.,  Electro  analysis,  1999,11, 546.

16- Matakova,R.N.,  Kenzhekhanova, Zh.T., and Badovamova,G.L.,  Jr.Anal

Chem.1999,54,1053.

17- Mediana –Escriche, J.,Sevillano– Cbeza, A. and  . Panella, M.A.M , Analyst, 1985,  110,807.

18- Pal, B.K., Mitra,B.K., Thakur, S. and  Pas, P.,  J.  Indian  Chem.  Soc.,  1988, 65, 862.

19- Abassi,S.A.,  Int. J.Enviorn. Anal. Chem.  1988 ,34,181;  Chem.Abs., 1989,111,32814v.

20- Rathore, D.P.S., and Tarafder, P.K.,  J.  Indian  Chem.  Soc.,  1991, 68, 179.

21- Kesava  Rao, C., Krishna  Reddy, V., and Sreenivasulu  Reddy, T.,  Talanta,  1994, 41, 237.

22- Agarmal, A., Chandra, S.K. and Pandex, L.P.,  Bull.  Mater.  Sci.,  1994,  17,  1467; Chem. Abs., 1995,123, 46912k.

23- Akama, Y., and  Kenjil, S., Res.  Bull.  Meisei.  Univ.  Phys.  Sci.  Eng.,  1994,  30,  15; Chem.  Abstr.  1995,  123, 217260c.

24-Saxena, R.C., Arana, H.C. and  Sachdeva, J.C.,  J.  Indian  Chem.  Soc.  1995, 72, 723.

25- Hassirozzaman, M., Sayed,K.A., Rahman, G.M.M., Rahmat, S.M.V., Alatag,H.M. and  Ismail,M.A., J.Bangladesh Chem. Soc.,1997,10, 25;  Chem.  Abs. 1999, 131, 345847b.

26- Yener, R., and Aycicek, A.,  Tr.  Fluoresc., 1998, 8,  45.

27- Yiquan, W. and Lan, G., Fenxi  Kexue  Xuebao, 2000,16,304; Chem.Abs. 2001,134, 12824v.

28- Masti SP, Seetharamappa J, Melwanki MB, Motahashi N,  Anal.Sci., 2002, 12(2), 167.

29- Goizman, M.S.,  Dokl.  Acad.  Nauk  USSR, 1964 ,184,599.

30- Kolthoff, I.M,Sandell, E.B.Mechan, E.J.and Bruckenstein,S., Quantitative Chemical  Analysis,  4th  Edition,1969 Collier– Macmillan Ltd., Library of Congress Cat.No.  69 – 10291.