Determination of ranitidine - HCl in pharmaceutical    formulations by kinetic spectrophotometric and flow injection - activated chemiluminescence  methods .

Abdul Majeed K . Ahamed , Ali I . Khaleel and Suham T. Amine

Chemistry Department, College of Education, Tikrit University.

(NJC)

# (Received on     3/1 /2006)          (Accepted for publication on   15/11/2006)

Abstract

This study involves development of  a simple  kinetic  spectrophotometric  method and  a new  flow injection-activated chemiluminescence (FIA-CL) for the determination of  ranitidine -HCl (R-HCl) in pharmaceutical preparations.

Spectrophotometric method was based  on the  oxidation  of the  R-HCl with alkaline potassium permanganate , the reaction is followed  spectrometrically by measuring the rate of  change of the absorbance at 600 nm. A fixed-time  (at 30min) method is adopted for  determining the drug concentration . The calibration graph was linear in the range of (1 - 7) μg-ml-1 with a correlation coefficient of 0.9981,detection limit of 0.183 μg-ml-1 and a relative standard deviation RSD% of 0.58-1.62%.

The method of FIA-CL was based on the activation of luminol-cobalt-H2O2 chemiluminescence by R-HCl. The linearity is (1-6) μg-ml-1 with detection  limit of (0.75) μg-ml-1 , and  correlation  coefficient was 0.9996 (n=6) and the relative  standard  deviation was (0.74-1.00)% .

The two methods were applied successfully to determine the content of       R-HCl in pharmaceutical preparations with a recovery of 97-99 % .

الخلاصة

تضمنت هذه الدراسة على جزئين أساسيين ، ضم الجزء الأول تطوير طريقة طيفية لتقدير دواء رانتدين هيدروكلورايد تستند على أكسدة هذا الدواء مع برمنغات البوتاسيوم القاعدية ( KMnO4 ) و متابعة التفاعل طيفياً من خلال قياس التغير بالامتصاص عند طول موجي 600 نانوميتر . وجد أن الزمن اللازم لتمامية التفاعل هو 30 دقيقة ومدى الخطية يتراوح بين ( 1 ـ 11 ) مايكرو غرام / مل ومعامل الارتباط 0.9981 ، وحد الكشف مقداره 0.183 مايكرو غرام / مل وقيم الانحراف القياسي النسبي RSD% تتراوح بين  (0.58 ـ 1.62)% .

أما الطريقة الثانية استندت على تحفيز البريق الكيميائي لنظام ( لومينول ـ بيروكسيد الهيدروجين ـ كوبلت ) بواسطة هذا الدواء وكانت الحدود الخطية للطريقة تترواح بين ( 1.0 ـ 10.0 ) مايكرو غرام / مل وبمعامل أرتباط 0.9996 وانحراف قياسي نسبي مئوي يتراوح بين 0.74% ـ 1.00% وحد كشــف ( 0.75 ) مايكرو غرام / مل .

طبقت هاتين الطريقتين بنجاح على المستحضرات الصيدلانية وكان الاسترداد المئوي (97ـ 99 )% .

Introduction

NO2

NHMe

.   HCl

Me2N

S

N

H

O

Ranitidine(1)- HCl ( R - HCl ):   N,N - Dim ethyl – 5 – [ 2 ( 1- methyl amino - 2 - nitro vinyl amino ) ethyl thio methyl ] furfuryl amine is a white or pale yellow , crystalline powder , freely soluble in water and methanol .

The empirical formula for R-HClH is C13H22N4O3S.HCl and its molecular weight is 350.9 , it is used as(2) inhibits gastric acid secretion  and its structure is:

This drug  has been determined in tablets and ampoules by  coluometric(3) and potentiometric(4)  titrations . HPTLC(5,6) and HPLC have been used for the determination of R-HCl using C18 column at different mobile phases(7-9) .

De Almeida(10) and chattoraj(11) used UV/Visible spectrophotometry for the determination of (R-HCl)in tablets pharmaceutical preparation, the absorbance was measured at 314nm with linearity ranged from 2-30 g/ml.

The present paper describes a spectrophotometric  method for the determination of (R-HCl )in pharmaceutical preparations based on the oxidation with alkaline potassium permanganate. This study also includes development of FIA-CL  method

Experimental

A) spectrophotometer apparatus for determination the absorbance, the spectrophotometer used was Philips model pu 8720 uv/visible, with quartus cells of 1 cm width.

B)  FIA Chemiluminescence apparatus

The FIA-CL configuration is outlined in figure 1 showing the system of chemical reaction used for the determination of (R-HCl).

Reagents

Reagents of analytical grade and distilled water were used  through out the study Solutions were prepared by appropriate dissolution as shown in table 1.

Luminol was prepared by dissolving in a solution of (0.1 M  Na2CO3) where as            KMnO4   was prepared and standardized with 0.1M Na2C2O4.

Hydrogen peroxide solution (1M) was prepared by diluting 45.72ml of H2O2 (48%) in 1 L of distilled water and standardized against standard 0.1 M KMnO4 .

0.55 ml of sulfuric acid (96%) was diluted in 100 ml  of distilled water and standardized against 0.1 M Na2CO3 . Cobalt (100μg.ml-1) was prepared  by dissolving  0.4039 gm of CoCl2.6H2O with 20 ml of 5 x10-3 M H2SO4  and  diluted to 1L  with  distilled water. Finally , R-HCl stock  solution (100μg.ml-1) was prepared by  dissolving 0.01 gm of  R-HCl powder in 1L of (0.7μg.ml-1) cobalt (II) solution . Solutions of lower concentrations were prepared by appropriate dilution .

Pharmaceutical preparation

Rantism tablets : provided from ( SDI ) Samara - Iraq .

Ten  tablets were grinded well and a certain portion of the final fine powder was accurately weighed to give an equivalent to 150 mg of ranitidine – HCl and then dissolved with (25ml) of Methanol . The resulting solution was washing by shaking with methanol and filtered on Whatman filter paper No.4 to remove any suspended particles . The filterate and washing were evaporated to dryness at 60C and the residue was redissolved in distilled water forming a solution of 100 μg.ml-1 concentration .

o

The same method was adopted for preparation of Zantac tablets, and the final dissolution was in (0.7 μg. ml-1) cobalt and ultrasonication was needed.

General procedure for the spectrophotometric method.

Initial rate method : Aliquots of 0.01M KMnO4 solution ( 1.0ml ) and 1.00M NaOH solution ( 2.0 ml ) were transferred into a 10ml volumetric flask. An accurate volume of the working solution of R- HCl (0.1 - 1.2) ml was added and diluted to volume with distilled water. The contents of the mixture were shaken well and immediately transferred to the spectrophotometric cell at room temperature. The absorbance was recorded at 600nm as a function of time against reagent blank.

In the second procedure, the absorbance was measured at a fixed time of (30min.) and was plotted against the final concentration of R-HCl and the content of the drug was calculated from either the calibration graph or regression equation .

Results and Discussion

Part(1 ) : Kinetic spectrophotometric determination of R- HCl in pharmaceutical preparation .

In an alkaline medium , potassium permanganate oxidizes R-HCl , resulting in the formation of manganate ion(13) , which showed an absorption peak at 600nm.

Because the intensity of the color increased with time , a kinetically based method was elaborated for the determination of R-HCl in dosage forms . Initial parameters for this study are given in table 2 .

The various experimental parameters affecting  the formation of the reaction product were optimised as follows :

Effect of the KMnO4 concentration

To study the effect of the KMnO4 concentration , aliquots of R-HCl containing 100μg .ml-1 were transferred into a series of 10ml volumetric flasks, followed by addition of varying volumes of 0.01M KMnO4 ( 0.1-1.1 ) ml and 2.0 ml of 1.00M NaOH solution . The absorbance at 600nm was measured at a fixed time of 30 minutes . It is apparent from figure 2 that the absorbance increased with increasing volume of the KMnO4 solution, and became constant at 0.9ml and so then 1.0ml of KMnO4 was used as the optimal volume .

Effect of the NaOH concentration

The influence of the NaOH concentration on the formation of MnO42- was examined critically. Figure3 shows that the maximum absorbance was obtained with 1.8ml of the 1.00M NaOH, so the optimum volume of 2.0ml was chosen.

Effect of oxidation time

It was found that the most acceptable value was obtained at a fixed time of 30min., and therefore was considered to be the most suitable time interval for measurements .

# Final absorbance spectrum of Ranitidine – HCl

A clear peak of MnO42- at 600nm ( fig.5 ) was obtained after the oxidation of R-HCl  by 0.01M KMnO4 solution in an alkaline medium of 1M NaOH .

The Ranitidine-HCl contains thio ether linkage which has an ability to be oxidized to the corresponding sulfoxide(13) . The reaction mechanism is proposed and given in the following equations .

+ 2MnO4- + 2 OH-

+ 2MnO4-2 + H2O

NO2

S

N

H

O

N

CH3

CH3

NHCH3

C

S

N

H

O

N

CH3

CH3

NHCH3

NO2

C

O

2

2

Recommended Analytical conditions

According to the results obtained , the optimum conditions for the determination of ranitidine-HCl using a kinetic spectrophotometric method are given in table 3 .

# Calibration graph

A linear calibration graph for ranitidine -HCl ( Fig.6 ) under the optimized conditions was obtained . Beers law is obeyed over the concentration range of 1.0- 7μg .ml-1 with correlation coefficient of 0.9981 and molar absorbance max 2.3 x 104 L.mol-1 .cm-1 .

The relative standard deviation of the method was 1.62 based on 6 replicate determinations. Table 4 shows summary of analytical data for the determination of ranitidine -HCl using spectrophotometric method .

## Pharmaceutical application

### The proposed method was applied for the determination of  ranitidine -HCl in tablets . Good precision and recovery were obtained . The method was successfully compared with the British pharmacopoeia standard method . The results obtained are summarized in table 5 .

Part (2) : Determination of ranitidine -HCl using a new FIA-CL method.

#### General procedure

The FIA-CL scheme is outlined in Figure1. Varying concentrations of R-HCl were prepared in Co2+ ( 0.7mg.ml-1) solution . 150ml aliquot of each solution was injected through the sample loop into the stream of Co2+ solution , which then was combined with luminol and hydrogen peroxide streams in the flow cell which situated in front of the photomultiplier tube (PMT) .

Results and discussion

The chemiluminescence of luminol–hydrogen peroxide–Co (II) System is very intense(14). However, in this work, trace amounts or R-HCl were found to be strongly activate  chemiluminescence signal of this system .

Effect of reagents concentration

The effect of luminol concentration  on the net chemiluminescence intensity was studied . Different concentrations of luminol (1×10-3 –1×10-5 M) are used to establish the best emission intensity  –  time profile that can be obtained. Figure 7 shows that (5 × 10-4 M) of luminol is the optimal  concentration .

The effect of H2O2 concentration was investigated ; from the results of

figure (8), the concentration of ( 1×10-2 M) H2O2  was selected to be the optimum concentration .

Effect of Sulphuric acid and Cobalt Concentrations

The effect of the acidity of cobalt Co2+ solution was also studied ; three concentration of cobalt Co2+  (0.2 , 0.5 , 0.7 ) μg . ml-1 with varing concentrations of (1 × 10-3 - 10 × 10-3 ) M of  H2SO4  were investigated . The best intensity was obtained at the concentration of (5x10-3)M of H2SO4(fig.9).

The effect of cobalt concentration was also studied, figure (10) shows that the suitable concentration of Co+2 is 0.7 μg / ml-1

It is worth noting here that as the concentration of H2SO4 increases the CL intensity increases , this is because of increasing the catalysed oxidation of luminol by H2O2 (16) . At 1x10-2 M H2SO4 the CL intensity decreases , this can be attributed to the cleavage of the formed fluorescent compound .

Table 6 shows that as the concentration of H2SO4 increases the time of analysis decreases accompanied with decreasing band width .

Effect of flow rate

A flow rate ranged from (1-10) ml/ min was investigated . Table 7 shows that the CL intensity increased with the increase of flow rate . However , a flow rate of 3ml/min is recommended for all streams because of satisfactory CL intensity , less  reagent consumption and short analysis time .

Recommended Analytical Conditions

A Summary of the optimum experimental conditions for the determination of R-HCl in pharmaceutical preparations is on table 8 .

Calibration graph

A calibration graph of relative chemiluminescence intensity against the ranitidine -HCl concentration was established by applying the optimal conditions . The regression equation is (Y = 7.7714 X + 0.1333 ) , and the linearity is in the range of (1.0 – 6 ) μg . ml-1 of  R-HCl ( fig. 11 ).The analytical data obtained from the calibration graph are summarized in  table 9 .

Interferences

Ranitidine - HCl is usually formulated in tablets , and injections forms , therefore, the effect of some common excipient substances usually present in pharmaceutical preparation were investigated . The presence of 1μg.ml-1 of Aerosil , and 2 μg.ml-1 of Mg-stearale gave no significant interfering effect on the chemiluminescence intensity of 5.0 μg.ml-1 of ranitidine -HCl .

Application of developed new FIA-CL method for the determination of R-HCl in pharmaceutical preparations

Two types of tablets containing R-HCl  were analysed using the developed method and the results were compared with the British pharmacopoeia standard method , Table 10 .

Comparison between the two methods

The two proposed methods were compared with other methods as shown in table (11). The value (0.55) of calculated F for FlA-CL / spectrophotometric methods is much less than the value ( 4.95) for tabulated F which indicates good agreement  .

Conclusion

In part (1) a simple kinetic spectrophotometric method was developed for the determination of ranitidine- HCl  in pharmaceutical preparation by oxidation of the R-HCl with alkaline KMnO4 . The results show good precision and accuracy for the determination of R-HCl in the range of (1-7) μg.ml-1 and correlation coefficient of 0.9981 with detection limit of 0.183 μg.ml-1 and RSD% of 1.62% with 99.0 % recovery .

In part (2)  the chemiluminescence was found to be activated by R-HCl and this is the base of the developed new method . The proposed method offers advantages of simplicity , rapidity , high sensitivity and low reagent consumption . The linearity of this method is (1-6) μg .ml-1 , correlation coefficient of 0.9996 (n=6) and RSD% was 1.00% with detection limit 0.75μg.ml-1and recovery (98%) .

Table (1) preparation of some solutions

| Substance | Molar Concentration (M) | Dissolved weight (gm) | Final Volume (ml) |
| --- | --- | --- | --- |
| ###### NaOH | 1.0 | 4.000 | 100 |
| NaCO3 | 0.1 | 10.599 | 1000 |
| Na2C2O4 | 0.1 | 13.390 | 1000 |
| KMnO4 | 0.1 | 15.800 | 1000 |
| Luminol | 1 x10-3 | 0.1772 | 1000 |

Table (2) : Initial Experimental parameters for the spectrophotometric method .

| Item | Preliminary parameter | Value |
| --- | --- | --- |
| 1 | Conc . of KMnO4 | 0.01 M |
| 2 | Conc . of NaOH | 1.00 M |
| 3 | Volume of KMnO4 | 1.0 ml |
| 4 | Volume of NaOH | 2.0 ml |
| 5 | Reaction Time | 6 min |
| 6 | Wave Length λmax (nm) | 600 nm |

Table (3): Optimum conditions for the detrmination of Ranitidine –HCl using the spectrophotometric method.

| No. | Parameter | Value |
| --- | --- | --- |
| 1. | Conc. Of KMnO4 | 0.01M |
| 2. | ##### Conc.of NaOH | 1.00 M |
| 3. | Fixed time | 30min |
| 4. | Vol . of KMnO4 | 1.0ml |
| 5. | Vol. of NaOH | 2.0 ml |
| 6. | Temperature | 25Cْ |

Table (4) : Analytical data for the determination of R-HCl by the spectrophotometric method.

| Analytical data | Value |
| --- | --- |
| Linear range | 1.0-7 μg. ml-1 |
| Correlation coefficient | 0.9981 |
| Regression equation | ###### Y=0.0461x +0.7324 |
| RSD % | 1.62 % |
| Detection limit | 0.183 μg. ml-1 |
| εmax | 2.3 x104 L.mol-1.Cm-1 |

Table (5) : Application of the proposed Spectrophotometric method for the determination of Ranitidine – HCl  in Pharmaceutical Preparations .

| Sample | Recovery % |  | RSD % |
| --- | --- | --- | --- |
|  | Proposed method * | Standard method |  |
| Pure R-HCl | 99.6 | 99.5 | 1.63 |
| Rantism Tablets | 97.8 | 97.3 | 1.62 |
| Zantac Tablets | 98.4 | 98.6 | 1.60 |

- Each result is the average of three determinations .

Table ( 6 ) Effect of sulphuric acid concentration on band width and time of analysis .

| H2SO4 Concentration (M) | Base band width (mm) | Time of analysis (sec) |
| --- | --- | --- |
| 1x10-4 | 26.0 | 75.0 |
| 5x10-4 | 22.0 | 67.0 |
| 1x10-3 | 16.0 | 55.0 |
| 5x10-3 | 10.0 | 45.0 |
| 1x10-2 | 8.5 | 43.0 |

Table (7) Effect of flow rate on the emission intensity , analysis time and peak width using 0.7 μg. ml-1 of Co+2.

| Flow rate ml/min | Intensity (mv) | Analysis time (Sec) | Peak width mm |
| --- | --- | --- | --- |
| 1.0 | 228 | 85 | 27 |
| 2.0 | 300 | 55 | 18 |
| 3.0 | 500 | 45 | 10 |
| 4.0 | 610 | 40 | 8 |
| 5.0 | 700 | 35 | 7 |
| 6.0 | 790 | 30 | 6 |
| 7.0 | 850 | 26 | 5 |
| 8.0 | 900 | 20 | 4 |
| 9.0 | 910 | 15 | 3.5 |
| 10.0 | 925 | 12 | 3.0 |

Table (8) The recommended analytical conditions for the determination of Ranitidine –HCl using FIA – Chemiluminescence system.

| Parameter | Recommended value |
| --- | --- |
| Conc. Of luminal | 5x10-4 M |
| Conc. Of H2O2 | 1x10-2 M |
| Conc. Of Co+2 | 0.7 μg. ml-1 |
| Conc. Of H2SO4 | 5x10-3 M |
| Flow rate | 3 ml/ min |
| Volume of Co+2 injected | 150 μl |

Table (9) Analytical data for the determination of Ranitidine –HCl using FIA-CL method.

| Analytical data | Value |
| --- | --- |
| Linear range | 1.0-6 μg. ml-1 |
| Correlation coefficient | 0.9996 |
| Regression equation | Y = 7.7714 X + 0.1333 |
| RSD % | 1.0% |
| Average Recovery | 99.0 % |
| Detection limit (D.L) | 0.75 μg. ml-1 |

Table (10) Application of the proposed method for the determination of Ranitidine –HCl in pharmaceutical preparations

| Sample | Recovery % |  | RSD % |
| --- | --- | --- | --- |
|  | Proposed method | Standard method |  |
| Pure Ranitidine-HCl | 99.0 | 99.3 | 1.76 |
| Rantism Tablets | 98.65 | 99.2 | 1.72 |
| Zantac Tablets | 98.30 | 99.12 | 1.69 |

Table (11) The statistical comparison of results for the spectrophotometric and FIA-CL method with other methods

| The method | Regression equation | Linearity (μ g/ml) | Correlation Coefficient (r) | Recovery % | RSD% | t-table | Ref. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ##### spectrophotometric | Y=0.0461 X+0.7321 | 1.0-7.0 | 0.9981 | 97-98 | 1.62 | 32.72 | ___ |
| FIA-CL | Y= 7.7714 X+0.1333 | 1.0-6.0 | 0.9996 | 98-99 | 1.00 | 63.26 | ___ |
| HPLC | _____ | 1.5-20 | 0.9978 | 93.00 | 4.50 | _____ | 5 |
| HPTLC | _____ | 5.0-10 | 0.9969 | 94.50 | 5.45 | _____ | 8 |
| Coulometric titration | _____ | 5.0-9.0 | _____ | 92.70 | 1.40 | _____ | 3 |
| Polarography | _____ | 3.5-15 | 0.9970 | 98.00 | 2.30 | _____ | 31 |

R

V

P

a

b

c

f

d

W

Luminol

H2O2

Water

Fig (1): Chemical reaction manifold used for the Flow-Injection Chemiluminescence determination of Ranitidine-HCl.

P- Peristaltic pump    V- Injection valve    F- Flow cell    d- Photomultiplier

R- Recorder    W- Waste solution .

0

0,05

0,1

0,15

0,2

0,25

0,3

0,35

0,4

0

0,2

0,4

0,6

0,8

1

1,2

Volume of KMnO4

(ml)

Absorbance

Fig (2) : Effect of the volume of 0.01M  KMnO4 on the intensity of color produced for the reaction (Ranitidine-HCl  10μg / ml , 2.0 ml of 1 M NaOH)

0,2

0,25

0,3

0,35

0,4

0,45

0,5

0,5

1

1,5

2

2,5

Volume of NaOH( ml)

Absorbance

Fig (3) : Effect of the volume of 1M NaOH on the intensity of color produced for

the reaction (Ranitidine-HCl 10 μg /ml , 1.0 ml of 0.01 M KMnO4)

Absorbance

Time , minutes

Fig (4) : Effect of time on the oxidation of R-HCl by 0.01M KMnO4

0

0.5

1

1.5

2

2.5

200

300

400

500

600

700

(A)

(B)

Absorbance

Wave length (nm)

Fig (5) : A-absorption spectrum of Ranitidine-HCl before oxidation

B- absorption spectrum of Ranitidine-HCl after oxidation.

Ranitidine- HCl μg/ ml

Absorbance

# Fig (6) : Calibration graph for the determenation of Ranitidine- HCl by the spectrophotometric method

0

100

200

300

400

500

600

2

4

6

8

10

12

Concentration of Luminol × 10-4 M

CL- Intensity (mv)

Fig (7)Effect of concentration of Luminol on the  CL- intensity.

0

100

200

300

500

600

0.5

1

1.5

2

Concentration of H2O2 ×10-2M

CL- Intensity (mv)

400

Fig (8) Effect of concentration of  H2O2  on the CL- intens.

100

150

200

250

300

350

400

450

500

550

600

2.5

5.0

7.5

10.0

Concentration Of H2SO4×10-3M

CL- Intensity  (mv)

Fig (9) Effect of Concentration of H2SO4 on the CL-Intensity .

0

100

200

300

400

500

600

0.2

0.4

0.6

0.8

1

1.2

Concentration of Cobalt Co+2 μg /ml-1

CL- Intensity (mv)

Figure (10) Effect of Cobalt concentration on the CL-intensity

Y = 7.7714 X + 0.1333

r =0.9996

10

20

30

40

50

0

2

4

6

8

Conc. Of  Ranitidine –HCl  µg/ml

HCl    μg /ml

CL intensity (mV)

Fig(11)calibration graph of Ranitidine –HCl by using FIA-CL method

References

- Martindale, the extra pharmacopoeia , the pharmaceutical press, London . (1993) 30th ed . pp.901 .

- Good man and Gilman’s , the pharmaceutical Basis of the rapeutics , MC Graw Hill , USA, (1996) . 9th ed . pp . 901.

- Nikolic , K .; Stankovic , B .; Bogavac , M. pharmazie , . 1995, 50 (4) . 301.

- Atkosar . Z. Tuncel , M. , Acta pharm . Turc ., 1989 , 31 (4) , 139 .

- Castro , A.; Ramero ,P. Gai , pharmazie, 2003 , 58 (10) , 696.

- Zendelovska , D.; Stafiov,T., Journal of pharmaceutical and Biomedical Analysis , 2003 , 33 (2) ,165.

- Ahmadiani , A .; Amini , H., J . of chromatography B : Biomedical Application , 2001 , 751 (2) , 291.

- Novakovic , J. J. of chromatography A , 1999 , 846 ( 1- 2 ) , 193.

- Mody , V.D .; Gandhi .; T.p.; Modi ,A. , Journal of chromatography B :

Biomedical Applications , 1996 , 676 (1) , 175.

- De Almeida Orsine , E.M.; Seferin Martins , J.L., Anal . Lett ., 1993 , 26

(9) , 1933.

- Chattaraj , S.C .;Das ,S.K .; Gupta ,B.K. , Indian Drugs , 1989 ,26 (7),

365.

- Barnett , N.W.; Hindson , B.J.; Lewis, S.W., Analytical chimica Acta , 1999 ,384(2) , 151.

- Rahman , N. , Kashif, M. , Analytical sciences , 2003 , 19 , 907.

- Shakir, I.M.A..; Azeez ,Y. J. and Al- Abachi , M. Q. , Iraq  J.Sci .,		1999 , 40A (3) , 122.

- Mahmoud, M.J., ph.D. Thesis , University of New Castle upon Tyne , (1981) .

- Martindale , the extra pharmacopoeia , the pharmaceutical press, London . (1993) 30th ed . pp.901 .

- Good man and Gilman’s , the pharmaceutical Basis of the rapeutics , MC Graw Hill , USA, (1996) . 9th ed . pp . 901.

- Nikolic , K .; Stankovic , B .; Bogavac , M. pharmazie ,  1995, 50 (4) . 301.

- Atkosar . Z. Tuncel , M. , Acta pharm . Turc ., 1989 , 31 (4) , 139 .

- Castro , A.; Ramero ,P. Gai , M. pharmazie, 2003 , 58 (10) , 696.

- Zendelovska , D.; Stafiov,T., Journal of pharmaceutical and Biomedical Analysis , 2003 , 33 (2) ,165.

- Ahmadiani , A .; Amini , H., J . of chromatography B : Biomedical Application , 2001 , 751 (2) , 291.

- Novakovic , J. J. of chromatography A , 1999 , 846 ( 1- 2 ) , 193.

- Mody , V.D .; Gandhi .; T.p.; Modi ,A. , Journal of chromatography B : Biomedical Applications , 1996 , 676 (1) , 175.

- De Almeida Orsine , E.M.; Seferin Martins , J.L., Anal . Lett ., 1993 , 26 (9) , 1933.

- Chattaraj , S.C .;Das ,S.K .; Gupta ,B.K. , Indian Drugs , 1989 ,26 (7) , 365.

- Barnett , N.W.; Hindson , B.J.; Lewis, S.W., Analytical chimica Acta , 1999 , 384(2) , 151.

- Rahman , N. , Kashif, M. , Analytical sciences , 2003 , 19 , 907.

- Shakir, I.M.A..; Azeez ,Y. J. and Al- Abachi , M. Q. , Iraq  J.Sci ., 1999 , 40A (3) , 122.

- Mahmoud, M.J., ph.D. Thesis , University of New Castle upon Tyne , (1981) .

- Richter, P.; Ines Toral, M.; Munoz Vargas ,F.; Analyst (London),  Jun 1994, 119(6), 1371.